UNIVERSITY OF GHANA

COLLEGE OF BASIC AND APPLIED SCIENCES
DEPARTMENT OF COMPUTER SCIENCE

SPIME: A STACKED PHYSICS-INFORMED META-ENSEMBLE FOR LOW-COST PM2.5 SENSOR CALIBRATION UNDER TROPICAL AND HARMATTAN METEOROLOGICAL REGIMES IN SUB-SAHARAN AFRICA

BY

NEAL ALLISON HUGHES
(STUDENT ID: 11178252)

A FINAL YEAR PROJECT REPORT SUBMITTED TO THE DEPARTMENT OF COMPUTER SCIENCE, UNIVERSITY OF GHANA, IN PARTIAL FULFILLMENT OF THE REQUIREMENTS FOR THE AWARD OF THE DEGREE OF
BACHELOR OF SCIENCE IN COMPUTER SCIENCE

SUPERVISOR:
PROFESSOR ISAAC WIAFE

JULY, 2026
LEGON, ACCRA, GHANA

DECLARATION

I hereby declare that this project report titled "SPIME: A Stacked Physics-Informed Meta-Ensemble for Low-Cost PM2.5 Sensor Calibration Under Tropical and Harmattan Meteorological Regimes in Sub-Saharan Africa" is the result of my own original research work conducted under the supervision of Professor Isaac Wiafe in the Department of Computer Science, University of Ghana, Legon. To the best of my knowledge, this work contains no material previously submitted for the award of any other degree or diploma at this or any other higher educational institution, except where due reference and academic acknowledgment have been explicitly made.

STUDENT:
Name: Neal Allison Hughes
Student ID: 11178252
Signature: ______________________                     Date: ______________________

SUPERVISOR CERTIFICATION:
I hereby certify that this project report was prepared by Neal Allison Hughes under my direct supervision, following the academic regulations for the award of the Bachelor of Science in Computer Science in the College of Basic and Applied Sciences, University of Ghana.

SUPERVISOR:
Name: Professor Isaac Wiafe
Department of Computer Science, University of Ghana
Signature: ______________________                     Date: ______________________

ABSTRACT

Context: Sub-Saharan Africa (SSA) faces a severe ambient air quality monitoring deficit due to the prohibitive capital and maintenance costs of regulatory-grade reference stations. While low-cost optical particle counters (such as the QuantAQ Modulair-PM) offer a scalable and cost-effective alternative for high-density spatial surveillance, their raw factory calibrations fail drastically under tropical West African conditions due to high-humidity hygroscopic aerosol swelling and heavy coarse mineral dust loadings during the dry Harmattan season.

Aim: To develop, evaluate, and deploy a physics-informed, non-overfitting machine learning calibration framework that corrects raw QuantAQ Modulair-PM sensor measurements to achieve reference-grade ambient PM2.5 monitoring accuracy against a co-located US EPA Federal Equivalent Method (FEM) Teledyne API T640 monitor in Accra, Ghana.

Method: Hourly co-located data (2024–2025) collected at the African Sensor Evaluation and Training Facility (Afri-SET), University of Ghana, Legon, was preprocessed and transformed into a 75-dimensional feature space. This feature vector incorporated multi-gamma empirical Köhler hygroscopic swelling curves fRH,γ, thermodynamic moisture parameters (Absolute Humidity, Vapor Pressure Deficit, Dew Point depression), multi-window EWMA time-series trends (1h–24h), circular temporal harmonics, and seasonal Harmattan dust indicators. A Stacked Physics-Informed Meta-Ensemble (SPIME v4) was developed using 5-Fold out-of-fold cross-validation across four diverse base learners (Random Forest, ExtraTrees, LightGBM, HistGradientBoosting) stacked through an L2-regularized RidgeCV meta-regressor, validated on an unseen 6-month out-of-sample holdout dataset (3,468 hourly records).

Result: Uncalibrated factory sensor outputs failed completely (R2=-0.0133, MAE=5.30 μg/m3), while baseline linear regression (R2=0.6760) and deep neural networks (R2≈0.49-0.52) exhibited poor accuracy and noise memorization. The proposed SPIME v4 model achieved state-of-the-art reference-grade performance on the unseen holdout dataset with R2=0.9632 (R2=0.9732 on test split), MAE=0.61 μg/m3 (a 66.3% error reduction over previous models), RMSE=1.23 μg/m3, and a near-zero generalization overfitting gap (-0.0002). Feature attribution confirmed that Köhler growth formulations and absolute humidity accounted for over 35% of model predictive power.

Conclusion: Integrating empirical aerosol hygroscopic physics with out-of-fold meta-ensemble stacking effectively eliminates sensor degradation and overfitting across extreme monsoon humidity and Saharan dust conditions. The resulting SPIME v4 pipeline delivers regulatory-grade PM2.5 monitoring capability, establishing a reliable, low-cost solution for air quality surveillance and environmental policy in Sub-Saharan Africa.

TABLE OF CONTENTS

Declaration	ii

Abstract	iii

List of Figures	vi

List of Tables	vii

List of Abbreviations	viii

CHAPTER 1: INTRODUCTION	1

1.1 Background and Motivation	1

1.2 Air Quality Challenges in Sub-Saharan Africa and West Africa	2

1.3 The Afri-SET Initiative and the Modulair-PM Project	3

1.4 Problem Statement	4

1.5 Research Aim and Objectives	5

1.6 Research Contributions	5

1.7 Scope of the Study	6

1.8 Organization of the Study	6

CHAPTER 2: LITERATURE REVIEW	8

2.1 Fine Particulate Matter (PM2.5) and Health Impacts in SSA	8

2.2 Low-Cost Sensors vs. Reference-Grade Monitors	9

2.2.1 Optical Particle Counters (OPCs) in LCS	9

2.2.2 Reference-Grade Monitors: Teledyne API T640	11

2.3 Environmental, Thermodynamic, and Meteorological Interferences	12

2.3.1 Relative Humidity (RH) and Aerosol Hygroscopicity	12

2.3.2 Temperature and Semi-Volatile Evaporation	14

2.3.3 The Harmattan Wind and Saharan Mineral Dust Dynamics	14

2.4 Sensor Calibration Paradigms and Machine Learning	16

2.4.1 Classical Linear Regression Models	16

2.4.2 Non-Linear Machine Learning Regressors	17

2.4.3 Multi-Layer Perceptrons and Deep Neural Networks	18

2.5 Time-Series Validation Strategies and Data Leakage Prevention	19

2.6 Physics-Informed Machine Learning and Stacking Meta-Ensembles	20

CHAPTER 3: RESEARCH METHODOLOGY	23

3.1 Experimental Co-location Site and Environmental Profile	23

3.2 Instrumentation and Sensor Specifications	24

3.2.1 Candidate LCS: QuantAQ Modulair-PM (MOD-PM-01068)	24

3.2.2 Reference Monitor: Teledyne API T640 FEM	25

3.3 Data Preprocessing and Quality Assurance Pipeline	26

3.3.1 Temporal Aggregation and Index Reindexing	26

3.3.2 Outlier Filtering and Quality Control	27

3.4 75-Dimensional Physics-Informed Feature Engineering	29

3.4.1 Multi-Gamma Köhler Hygroscopic Growth Formulations	29

3.4.2 Atmospheric Thermodynamics (VPD, AH, Dew Point)	31

3.4.3 Multi-Window EWMA and Rolling Trend Features	32

3.4.4 Cyclical Temporal Harmonic Encodings	33

3.4.5 Harmattan Seasonal Dust Indicators and Scoring	33

3.5 Model Architectures and Calibration Paradigms	35

3.5.1 Baseline: Multiple Linear Regression (MLR)	35

3.5.2 Bagged Ensemble: Random Forest Regressor	35

3.5.3 Deep Learning: Multi-Layer Perceptron (MLP)	36

3.5.4 Primary Model: Stacked Physics-Informed Meta-Ensemble (SPIME v4)	36

3.6 Model Validation and Out-of-Sample Testing Strategy	39

3.7 Performance Evaluation Metrics and Residual Diagnostics	40

CHAPTER 4: EXPERIMENTAL RESULTS AND DISCUSSION	43

4.1 Comparative Model Calibration Performance	43

4.2 Diagnostic Analysis of Overfitting across Paradigms	46

4.2.1 Failure of Factory Default Sensor Calibrations	46

4.2.2 Neural Network Noise Memorization and Generalization Deficit	47

4.2.3 Tree Ensemble Overfitting vs. SPIME v4 Breakthrough	48

4.3 Physical Interpretation of Köhler Swelling and Feature Attribution	49

4.4 Seasonal Performance: Monsoon Wet Season vs. Harmattan Dry Season	52

4.5 Residual Error Diagnostics and Statistical Verification	54

CHAPTER 5: CONCLUDING REMARKS	57

5.1 Summary of Key Findings	57

5.2 Summary of Research Contributions	58

5.3 Practical Deployment Recommendations for Afri-SET	59

5.4 Limitations of the Study	60

5.5 Directions for Future Research	61

REFERENCES / BIBLIOGRAPHY	63

APPENDIX A: FEATURE ENGINEERING PIPELINE IMPLEMENTATION	67

APPENDIX B: HYPERPARAMETER TUNING CONFIGURATIONS	69

LIST OF FIGURES

Figure 2.1: Optical Particle Counter (OPC) Operating Principle based on Mie Light Scattering Theory	10

Figure 2.2: Aerosol Hygroscopic Particle Growth and Deliquescence Phase Transition Curve	13

Figure 2.3: Seasonal Meteorological Regimes and Aerosol Dynamics in West Africa: Harmattan vs. Monsoon	15

Figure 3.1: Geographical Location of Afri-SET Co-location Facility at University of Ghana, Legon	24

Figure 3.2: End-to-End Machine Learning Calibration Pipeline Architecture	26

Figure 3.3: Monthly Scatter Distribution of Raw Modulair-PM PM2.5 vs. Teledyne T640 Reference Measurements	28

Figure 3.4: Complete Hourly Time-Series Tracking of Raw Sensor PM2.5 vs. Reference Teledyne T640 Ground Truth	28

Figure 3.5: Multi-Gamma Empirical Köhler Hygroscopic Aerosol Growth Curves across Exponents γ ∈ [0.05, 0.50]	30

Figure 3.6: Stacked Physics-Informed Meta-Ensemble (SPIME v4) Pipeline Architecture	37

Figure 3.7: Chronological Train-Test and Unseen 6-Month Out-of-Sample Holdout Partitioning Scheme	39

Figure 4.1: Scatter Plot of Actual Teledyne T640 Reference PM2.5 vs. Calibrated SPIME v4 Predictions on Unseen Holdout Data	44

Figure 4.2: High-Resolution Time-Series Tracking of Calibrated SPIME v4 PM2.5 Predictions against Teledyne Reference Data (300-Hour Subset)	45

Figure 4.3: Comparative Time-Series Tracking: Multi-Layer Perceptron (MLP) Neural Network vs. Ground Truth Reference Data	47

Figure 4.4: Top 20 Feature Importance Rankings Extracted from SPIME Base Estimators highlighting Köhler Physics and Thermodynamic Predictors	50

Figure 4.5: Residual Error Histogram and Normal Density Fit for SPIME v4 on Unseen Holdout Data	54

Figure 4.6: Normal Quantile-Quantile (Q-Q) Diagnostic Plot of SPIME v4 Standardized Residuals	55

Figure 4.7: Homoscedasticity Diagnostic: Residual Errors vs. Calibrated Predicted PM2.5 Concentrations	56

LIST OF TABLES

Table 2.1: Comparison of Regulatory FEM Particulate Monitors and Low-Cost Optical Particle Counters	11

Table 2.2: Systematic Summary of Machine Learning Calibration Literature for Low-Cost Air Quality Sensors (2017–2025)	21

Table 3.1: Technical Specifications of the QuantAQ Modulair-PM Sensor (MOD-PM-01068)	25

Table 3.2: Technical Specifications of the Teledyne API T640 Reference Monitor	25

Table 3.3: Systematic Summary of the 75 Engineered Physics-Informed Atmospheric Features	34

Table 3.4: Optimized Hyperparameter Configurations for SPIME Level 0 Base Estimators and Level 1 Meta-Learner	38

Table 4.1: Comparative Performance Metrics on Out-of-Sample Holdout Data (3,468 Unseen Hourly Samples)	44

Table 4.2: Seasonal Performance Breakdown: Monsoon Wet Season vs. Harmattan Dry Season on Unseen Holdout Data	52

Table 4.3: Generalization and Overfitting Gap Diagnostics across Model Architectures	48

Table B.1: Hyperparameter Search Spaces and Selected Optimal Parameter Configurations	69

LIST OF ABBREVIATIONS

Table 1: Alphabetical List of Acronyms and Technical Abbreviations

| Abbreviation | Complete Expanded Meaning / Description |
| --- | --- |
| Afri-SET | African Sensor Evaluation and Training Facility |
| AH | Absolute Humidity (g/m³) |
| AI | Artificial Intelligence |
| API | Application Programming Interface / Teledyne Advanced Pollution Instrumentation |
| ASC | Aerosol Sample Conditioner |
| BAM | Beta Attenuation Monitor |
| CO | Carbon Monoxide |
| COPD | Chronic Obstructive Pulmonary Disease |
| CV | Cross-Validation |
| DOE | Department of Energy |
| DRH | Deliquescence Relative Humidity |
| EPA | Environmental Protection Agency |
| ET | ExtraTrees Regressor |
| EWMA | Exponentially Weighted Moving Average |
| FEM | Federal Equivalent Method |
| FRM | Federal Reference Method |
| Gini | Gini Impurity Index |
| HGB | HistGradientBoosting Regressor |
| IQR | Interquartile Range |
| LCS | Low-Cost Sensor |
| LED | Light Emitting Diode |
| LGB / LightGBM | Light Gradient Boosting Machine |
| LLM | Large Language Model |
| LPM | Liters Per Minute |
| MAE | Mean Absolute Error (µg/m³) |
| MBE | Mean Bias Error (µg/m³) |
| ML | Machine Learning |
| MLP | Multi-Layer Perceptron (Neural Network) |
| MLR | Multiple Linear Regression |
| MoE | Atmospheric Mixture of Experts |
| NO2 | Nitrogen Dioxide |
| O3 | Ozone |
| OLS | Ordinary Least Squares |
| OOF | Out-of-Fold (Cross-Validation Stacking) |
| OPC | Optical Particle Counter |
| PAH | Polycyclic Aromatic Hydrocarbons |
| PM | Particulate Matter |
| PM1 | Particulate Matter ≤ 1.0 µm aerodynamic diameter |
| PM2.5 | Particulate Matter ≤ 2.5 µm aerodynamic diameter |
| PM10 | Particulate Matter ≤ 10.0 µm aerodynamic diameter |
| Q-Q Plot | Quantile-Quantile Probability Diagnostic Plot |
| R² | Coefficient of Determination |
| RF | Random Forest Regressor |
| RH | Relative Humidity (%) |
| RidgeCV | L2-Regularized Ridge Regression with Cross-Validation |
| RMSE | Root Mean Squared Error (µg/m³) |
| SLR | Simple Linear Regression |
| SPIME | Stacked Physics-Informed Meta-Ensemble |
| SSA | Sub-Saharan Africa |
| SVOC | Semi-Volatile Organic Compounds |
| UTC | Coordinated Universal Time |
| VPD | Vapor Pressure Deficit (kPa) |
| WHO | World Health Organization |
| WMO | World Meteorological Organization |

CHAPTER 1: INTRODUCTION

1.1 Background and Motivation

Ambient air pollution constitutes one of the most severe public health crises of the twenty-first century, contributing to an estimated 6.7 million premature deaths globally each year (World Health Organization, 2021). Among various airborne contaminants, fine particulate matter with an aerodynamic diameter equal to or less than 2.5 micrometers (PM2.5) represents the most hazardous pollutant to human physiology. Because of its microscopic aerodynamic dimensions, PM2.5 can bypass the upper respiratory defenses of the nasal cavity, penetrate deeply into the pulmonary alveoli, and translocate directly into the systemic bloodstream (Mauderly & Chow, 2008). Long-term inhalation of elevated PM2.5 concentrations is epidemiologically linked to ischemic heart disease, stroke, chronic obstructive pulmonary disease (COPD), lung carcinoma, and acute lower respiratory infections in young children.

Despite the acute public health risks associated with ambient air pollution, atmospheric monitoring infrastructure remains critically sparse globally, with the most severe data deserts concentrated across low- and middle-income regions. Sub-Saharan Africa (SSA), which contains some of the world's most rapidly urbanizing populations, suffers from an alarming lack of continuous, ground-level air quality monitoring stations (DeSouza et al., 2017). Establishing and maintaining conventional regulatory reference stations—such as Beta Attenuation Monitors (BAM) or broad-spectrum optical spectroscopy systems like the Teledyne API T640—demands prohibitive capital expenditures exceeding 20,000 to 35,000 per unit, coupled with substantial ongoing operational costs, continuous climate-controlled enclosures, highly trained technical personnel, and uninterrupted grid electricity. Consequently, vast municipal populations across West Africa live and work in environments without any active, real-time air quality surveillance.

In response to this severe monitoring deficit, low-cost air quality sensors (LCS) have emerged as a transformative technology capable of democratizing atmospheric data collection (Clements et al., 2017; Rai et al., 2019). These low-cost instruments, which typically rely on optical particle counters (OPCs) such as the Alphasense OPC-N3, are compact, power-efficient, and cost a fraction of regulatory-grade stations (typically 200 to 2,000 per unit). Dense spatial deployments of low-cost sensors enable municipal planners, atmospheric scientists, and public health authorities to map intra-urban pollution gradients at unprecedented spatial and temporal resolutions, uncovering localized emissions hotspots and informing targeted environmental policy (Hagan et al., 2019).

However, the operational deployment of low-cost optical sensors is severely compromised by profound data quality vulnerabilities. Low-cost optical particle counters do not measure particulate mass directly; instead, they count light scattering pulses as individual particles pass through a focused laser beam, estimating mass concentrations via Mie scattering theory under idealized assumptions regarding particle spherical geometry, density, and refractive index (Crilley et al., 2018). In real-world ambient atmospheres, these mathematical assumptions break down due to three primary environmental interferences:

Hygroscopic Aerosol Growth under Elevated Relative Humidity: Tropical ambient moisture causes soluble aerosol particles to absorb water and swell in size once relative humidity exceeds the Deliquescence Relative Humidity (DRH) threshold (70%–80%). Optical sensors interpret these water-swollen droplets as larger solid particles, drastically overestimating dry particulate mass concentrations (Di Antonio et al., 2018; Pope et al., 2018).

Aerosol Composition Transitions during Saharan Harmattan Winds: Seasonal meteorological shifts, specifically the dry, dust-laden Harmattan trade winds originating from the Sahara Desert, transport immense loadings of coarse, highly non-spherical, high-density mineral dust particles. These mineral aerosols alter optical scattering properties and cause severe coincidence errors, leading to substantial measurement divergence (Sun et al., 2019; Westervelt et al., 2019).

Physical Sensor Aging and Baseline Drift: Extended field deployment leads to optical degradation, photodiode sensitivity decay, dust accumulation on optical flow chambers, and sampling fan degradation, causing gradual baseline calibration drift (Feenstra et al., 2019).

To overcome these environmental vulnerabilities, localized, season-aware machine learning calibration models are required. Machine learning algorithms can capture complex, non-linear interactions between raw optical scattering signals, ambient temperature, relative humidity, and seasonal dust events, correcting low-cost sensor data to reference-grade standards (Zimmerman et al., 2018; Malings et al., 2019; Lim et al., 2021).

1.2 Air Quality Challenges in Sub-Saharan Africa and West Africa

The West African sub-region, and specifically coastal metropolitan centers such as Accra, Ghana, presents a highly unique and challenging atmospheric environment for particulate matter monitoring. The region is characterized by a complex convergence of diverse anthropogenic emissions and powerful natural meteorological drivers (Appoh et al., 2020):

Intense Anthropogenic Emissions: Rapid population expansion and uncontrolled urban development have resulted in high traffic density with aged, poorly maintained diesel fleets, extensive reliance on solid biomass fuels (wood and charcoal) for domestic cooking, ubiquitous open refuse combustion, and unpaved road dust resuspension.

Seasonal Saharan Mineral Dust Incursions: Between November and March, coastal West Africa is dominated by the Harmattan winds—northeasterly trade winds that carry massive plumes of mineral dust from the Bodélé Depression and Saharan basin. During intense Harmattan dust episodes, ambient PM2.5 concentrations routinely exceed 150 to 300 µg/m³, while relative humidity drops below 25% (Sun et al., 2019; World Meteorological Organization, 2021).

Tropical Monsoon Humidity and Washout Dynamics: From April to October, the West African Monsoon brings persistent tropical convective rainfall and elevated relative humidity, frequently exceeding 85% to 95%. While intense precipitation events induce atmospheric washout—rapidly scavenging suspended particulates—the sustained high humidity causes extreme hygroscopic particle swelling.

Standard factory calibration algorithms supplied by low-cost sensor manufacturers are developed in temperate European or North American laboratory chambers under low-humidity conditions and combustion-dominated aerosol regimes. When these factory models are deployed in the high-humidity, high-dust conditions of Accra, they fail catastrophically, exhibiting negative correlation coefficients (R² < 0) and systematic measurement biases exceeding 100% (Westervelt et al., 2019; Maclean et al., 2021).

1.3 The Afri-SET Initiative and the Modulair-PM Project

This research is embedded within the experimental operations of the African Sensor Evaluation and Training Facility (Afri-SET), situated at the Department of Physics, University of Ghana, Legon, Accra. Afri-SET serves as the premier regional testing and calibration hub in West Africa, dedicated to evaluating candidate low-cost environmental sensing technologies under authentic Sub-Saharan meteorological and atmospheric conditions.

The core experimental platform at Afri-SET co-locates candidate low-cost monitoring units with US EPA Federal Equivalent Method (FEM) reference-grade instrumentation. In this study, we investigate the QuantAQ Modulair-PM (sensor unit MOD-PM-01068), an advanced low-cost ambient air quality monitor incorporating an Alphasense OPC-N3 optical particle counter alongside precision relative humidity, temperature, and atmospheric pressure sensors. Co-located adjacent to a Teledyne API T640 continuous particulate spectrometer, a high-resolution hourly dataset spanning the years 2024 to 2025 was curated. This multi-season dataset captures the complete atmospheric spectrum of West Africa—from severe Saharan dust episodes to monsoon humidity surges—providing an empirical foundation for training physics-informed machine learning models.

1.4 Problem Statement

Low-cost optical particulate sensors (such as the QuantAQ Modulair-PM) offer the only financially viable mechanism to overcome the severe ambient air quality data deficit across Sub-Saharan Africa. However, raw sensor measurements suffer from extreme inaccuracies caused by non-linear hygroscopic aerosol swelling under tropical humidity, optical saturation during Saharan dust storms, and baseline sensor drift. Standard linear calibration models cannot capture the non-linear physics of Köhler droplet growth, while deep neural networks and unregularized tree ensembles frequently suffer from severe data leakage and overfitting on autocorrelated environmental time series. Furthermore, existing calibration studies in Africa have rarely incorporated multi-gamma Köhler growth formulations, thermodynamic vapor pressure deficit terms, or regional Harmattan dust indicators into a unified, non-overfitting meta-ensemble architecture. Without robust, physics-informed calibration pipelines, low-cost sensor networks cannot deliver the data fidelity required for regulatory compliance, epidemiological exposure assessment, or environmental policymaking in West Africa.

1.5 Research Aim and Objectives

Primary Aim: To develop, evaluate, and deploy a physics-informed, non-overfitting machine learning calibration framework (SPIME v4) that corrects raw QuantAQ Modulair-PM sensor measurements to achieve regulatory reference-grade PM2.5 monitoring accuracy against a co-located FEM Teledyne API T640 monitor in Accra, Ghana.

Specific Objectives:

Curate and Preprocess Co-located Database: Synchronize, reindex, and quality-filter hourly particulate and meteorological measurements from the QuantAQ Modulair-PM sensor and Teledyne T640 reference monitor collected at the University of Ghana, Legon across 2024–2025.

Formulate Physics-Informed Features: Engineer a comprehensive 75-dimensional feature space integrating multi-gamma empirical Köhler hygroscopic growth equations fRH,γ, thermodynamic water mass parameters (Absolute Humidity, Vapor Pressure Deficit, Dew Point depression), multi-window EWMA time-series trends (1h to 24h), cyclical temporal harmonics, and regional Harmattan dust storm indicators.

Construct Stacked Meta-Ensemble Architecture (SPIME v4): Develop a two-level calibration architecture combining four diverse base estimators (Random Forest, ExtraTrees, LightGBM, HistGradientBoosting) with 5-Fold out-of-fold cross-validated stacking and an L2-regularized RidgeCV meta-regressor.

Evaluate Generalization and Overfitting Elimination: Validate calibration performance across chronological partitions and an unseen 6-month out-of-sample holdout dataset (3,468 hourly records), assessing R2, MAE, RMSE, systematic bias, and generalization gap (Rtrain2-Rtest2).

Conduct Residual Diagnostics and Feature Attribution: Perform comprehensive error diagnostics (normality, Q-Q plots, homoscedasticity) and extract tree-based feature importances to quantify the physical drivers of optical sensor error correction.

1.6 Research Contributions

This study provides several substantial contributions to environmental data science and low-cost sensor calibration in Sub-Saharan Africa:

First Integration of Multi-Gamma Köhler Physics in SSA Sensor Calibration: Formulated eight continuous empirical Köhler hygroscopic swelling functions that provide decision trees with explicit physical laws of aerosol water uptake, eliminating reliance on arbitrary humidity splits.

Development of the SPIME v4 Meta-Ensemble: Designed a stacked calibration architecture that leverages out-of-fold cross-validation stacking to achieve reference-grade accuracy (Holdout R2=0.9632, MAE=0.61 μg/m3) while completely eliminating generalization overfitting (gap = -0.0002).

Harmattan Dust Storm Adaptation: Introduced domain-specific Harmattan scoring algorithms that protect the calibration model against over-correcting dry mineral dust as humidity swell during the Saharan dust season.

Deployable Open-Source Pipeline for Afri-SET: Encapsulated the trained SPIME v4 model into a lightweight, deployable Python artifact for automated real-time calibration on the Afri-SET web portal.

1.7 Scope of the Study

This study is focused on the calibration of fine particulate matter (PM2.5) measurements from the QuantAQ Modulair-PM sensor (unit MOD-PM-01068) co-located at the Department of Physics, University of Ghana, Legon, Accra. Ground truth reference data is provided exclusively by the co-located Teledyne API T640 optical mass spectrometer. Environmental covariates are restricted to co-located relative humidity, ambient temperature, atmospheric pressure, and regional dust indicators. The study does not evaluate coarse particulates (PM10) or gaseous criteria pollutants (NO2, O3, CO), nor does it evaluate spatial sensor relocation outside Greater Accra.

1.8 Organization of the Study

The remainder of this project report is structured as follows:

Chapter 2: Literature Review examines the theoretical foundations of optical particle counters, Mie light scattering, aerosol hygroscopic swelling, Saharan dust transport, classical and machine learning calibration paradigms, time-series data leakage, and recent advances in physics-informed stacking ensembles.

Chapter 3: Research Methodology describes the experimental co-location facility at Afri-SET, sensor technical specifications, data preprocessing pipelines, the 75-dimensional feature engineering formulation, the SPIME v4 meta-ensemble architecture, validation partitions, and evaluation metrics.

Chapter 4: Experimental Results and Discussion presents empirical performance metrics on unseen holdout data, comparative evaluations across six model paradigms, overfitting diagnostics, feature importance rankings, seasonal performance breakdowns, and residual error tests.

Chapter 5: Concluding Remarks summarizes the major findings against the research objectives, details practical deployment guidelines for Afri-SET, identifies study limitations, and outlines directions for future research.

CHAPTER 2: LITERATURE REVIEW

2.1 Fine Particulate Matter (PM2.5) and Health Impacts in SSA

Atmospheric fine particulate matter with an aerodynamic diameter ≤2.5 μm (PM2.5) is recognized as the single largest environmental risk factor for premature mortality worldwide (World Health Organization, 2021). Due to their minute physical dimensions, PM2.5 aerosols remain suspended in ambient air for days to weeks and can be transported across continental distances by atmospheric winds. When inhaled, these fine particles penetrate past the ciliated epithelial cells of the upper trachea, reaching the deep gas-exchange regions of the pulmonary alveoli (Mauderly & Chow, 2008). Insoluble fractions translocate through the alveolar-capillary membrane into systemic circulation, inducing oxidative stress, vascular endothelial dysfunction, pulmonary inflammation, and systemic atherosclerosis.

In Sub-Saharan Africa (SSA), the public health burden of fine particulate matter is exacerbated by the coexistence of intense local combustion sources and immense natural aerosol emissions (DeSouza et al., 2017). Rapid urban population growth has outpaced municipal infrastructure development, leading to dense informal settlements where solid biomass (firewood and charcoal) remains the primary energy source for domestic cooking and heating. Simultaneously, high densities of aged, second-hand commercial diesel vehicles, ubiquitous two-stroke motorcycle taxis, and open burning of municipal solid waste generate heavy concentrations of carbonaceous soot, polycyclic aromatic hydrocarbons (PAHs), and toxic heavy metals (Appoh et al., 2020).

Despite this high pollution burden, regulatory-grade air quality surveillance across Sub-Saharan Africa remains critically inadequate. In 2021, fewer than 10% of countries in Africa had active, publicly accessible continuous particulate monitoring networks (World Health Organization, 2021). This data deficit severely hinders public health epidemiological research, prevents municipal authorities from issuing timely air quality advisories, and leaves environmental protection agencies without empirical data to formulate evidence-based clean air regulations.

2.2 Low-Cost Sensors vs. Reference-Grade Monitors

To understand the physical origins of low-cost sensor measurement errors, it is necessary to contrast their optical sensing mechanisms with regulatory reference-grade instrumentation.

2.2.1 Optical Particle Counters (OPCs) in LCS

Low-cost particulate matter monitors, such as the QuantAQ Modulair-PM (which integrates the Alphasense OPC-N3), rely on optical particle counter (OPC) technology (Crilley et al., 2018). Inside the device, a miniature internal sampling fan draws ambient air through a light-shielded optical chamber. A focused continuous-wave laser diode (typically operating at a wavelength λ=658 nm) illuminates the incoming aerosol stream. When an individual particle intersects the laser focal volume, it scatters light in all directions, as illustrated in Figure 2.1.

Figure 2.1: Optical Particle Counter (OPC) Operating Principle based on Mie Light Scattering Theory

A high-speed silicon photodiode detector positioned at a specific scattering angle captures the scattered light pulse. The peak amplitude and pulse area of the detected electrical signal are proportional to the scattering cross-section of the particle, governed by Mie Scattering Theory (Crilley et al., 2018). Under Mie theory, the intensity of light scattered by a spherical particle depends on its physical diameter (D_p), laser wavelength (λ), scattering angle (θ), and the complex refractive index of the aerosol material m=n-i⋅k:

I(θ, D_p, m) = (I_0 / (2 k_w² r²)) · [ |S_1(θ, x, m)|² + |S_2(θ, x, m)|² ]	(Eq. 2.1)

where x=π⋅Dp/λ is the dimensionless size parameter, kw=2π/λ is the optical wavenumber, and S_1, S_2 are the complex Mie scattering amplitude functions. The OPC sorts detected electrical pulses into discrete size bins (e.g., 24 bins spanning 0.35 µm to 40 µm in the OPC-N3). To estimate particulate mass concentration (µg/m³), the sensor's firmware computes:

PM_mass = ∑ [ N_i · V_i · ρ_i ] = ∑ [ N_i · (π / 6) · D_i³ · ρ_i ]	(Eq. 2.2)

where N_i is the particle count in bin i, D_i is the geometric mean diameter of bin i, and ρi is the assumed aerosol particle density.

This mathematical conversion introduces severe systematic errors in ambient deployments because:

Refractive Index Heterogeneity: The factory algorithm assumes a constant refractive index (typically m = 1.5 - 0.0i), whereas ambient aerosols are complex mixtures of light-absorbing black carbon (m≈1.85-0.71i), non-absorbing sulfates (m≈1.53-0.0i), and irregularly shaped mineral dust (m≈1.56-0.005i) (Hagan et al., 2019).

Particle Density Variations: A uniform density of ρ=1.65 g/cm3 is assumed, whereas combustion soot has an effective density of 1.0-1.2 g/cm3 and Saharan mineral dust has a density of 2.5-2.7 g/cm3.

Non-Spherical Particle Geometry: Mineral dust particles are highly angular and irregular, violating the spherical volume assumption (V=π/6⋅D3) and causing substantial particle sizing errors (Sun et al., 2019).

2.2.2 Reference-Grade Monitors: Teledyne API T640

In contrast to low-cost sensors, regulatory reference monitors utilize sophisticated optical and physical measurement systems. The Teledyne API T640 is a continuous particulate monitor designated as a US EPA Federal Equivalent Method (FEM EQPM-0514-217) for PM2.5 monitoring. The T640 employs broad-spectrum white LED light-scattering polychromatic spectroscopy. By utilizing a polychromatic white light source across multiple scattering angles, the T640 eliminates the monochromatic Mie resonance oscillations that plague single-wavelength lasers, providing highly accurate particle sizing independent of aerosol refractive index. Furthermore, the T640 incorporates an automated heated inlet (Aerosol Sample Conditioner) that dynamically removes excess moisture without volatilizing semi-volatile organic compounds. Table 2.1 summarizes the architectural differences between reference monitors and low-cost sensors.

Table 2.1: Comparison of Regulatory FEM Particulate Monitors and Low-Cost Optical Particle Counters

| Feature / Parameter | Teledyne API T640 (FEM Reference) | QuantAQ Modulair-PM (Alphasense OPC-N3) |
| --- | --- | --- |
| Optical Light Source | Broad-Spectrum Polychromatic White LED | Monochromatic Laser Diode (λ = 658 nm) |
| Measurement Principle | Polychromatic 90° Light Scattering Spectroscopy | Single-Angle Mie Light Pulse Sizing |
| Moisture Conditioning | Dynamic Heated Aerosol Conditioner (Active Inlet) | Unheated Ambient Flow Chamber (Passive) |
| Particle Sizing Range | 0.18 µm to 20.0 µm (64 channels) | 0.35 µm to 40.0 µm (24 discrete bins) |
| Mass Concentration Range | 0.1 to 10,000 µg/m³ | 0.1 to 2,000 µg/m³ |
| Regulatory Status | US EPA Federal Equivalent Method (FEM) | Non-regulatory Screening Instrument |
| Capital Expenditure | 25,000 – 35,000 USD | 1,500 – 2,500 USD |
| Power & Infrastructure | High Power (150W), Climate-Controlled Shelter | Low Power (< 5W), Solar / Battery Operable |

2.3 Environmental, Thermodynamic, and Meteorological Interferences

Meteorological covariates exert a profound non-linear influence on uncalibrated low-cost optical sensors, acting as the primary source of measurement divergence from reference ground truth.

2.3.1 Relative Humidity (RH) and Aerosol Hygroscopicity

The most pervasive environmental interference in tropical environments is aerosol hygroscopicity—the phenomenon whereby atmospheric aerosols absorb ambient water vapor and undergo physical swelling (Di Antonio et al., 2018; Pope et al., 2018). As relative humidity increases, aerosol particles experience a phase transition known as deliquescence at the Deliquescence Relative Humidity (DRH) threshold (typically 70%–80% RH for common ammonium sulfate and organic aerosol mixtures), transforming from solid cores into aqueous droplets, as illustrated in Figure 2.2.

Figure 2.2: Aerosol Hygroscopic Particle Growth and Deliquescence Phase Transition Curve

Under classical Köhler Theory (Köhler, 1936), the equilibrium diameter (D_p) of a water-swollen aerosol droplet is governed by the water vapor saturation ratio (S=RH/100):

ln(RH / 100) = (A / D_p) - (B / D_p³)	(Eq. 2.3)

where A represents the Kelvin curvature effect (which increases vapor pressure over curved droplet surfaces) and B represents the Raoult solute effect (which decreases vapor pressure due to dissolved salts). In low-cost sensor calibration literature, empirical hygroscopic mass growth is commonly parameterized using the f(RH) enhancement factor (Crilley et al., 2018; Di Antonio et al., 2018):

fRH=PMmeasured/PMdry=1+κ·RH/100/1-RH/100γ	(Eq. 2.4)

where κ is the bulk hygroscopicity parameter and γ is an empirical growth exponent. Because uncalibrated OPCs detect swollen droplet diameters, they calculate an inflated particulate volume (proportional to Dp3), overestimating true dry mass by 50% to over 200% under high monsoon humidity (RH > 80%).

2.3.2 Temperature and Semi-Volatile Evaporation

Ambient temperature fluctuations influence both aerosol thermodynamics and internal sensor electronics. Elevated ambient temperatures (> 32°C) induce the evaporation of semi-volatile organic compounds (SVOCs) and ammonium nitrate from the aerosol phase into the gas phase, reducing particulate mass on reference monitors while altering the refractive index of remaining aerosol particles (Malings et al., 2019). Furthermore, temperature changes affect laser diode emission intensity, silicon photodiode dark current, and internal amplifier gain, requiring explicit thermodynamic temperature compensation.

2.3.3 The Harmattan Wind and Saharan Mineral Dust Dynamics

In coastal West Africa, the dry season (November to March) is dominated by the Harmattan—a continental northeasterly trade wind system that transports immense quantities of mineral dust from the Bodélé Depression and Sahara Desert southwestward across the Gulf of Guinea (Sun et al., 2019; Appoh et al., 2020), as conceptualized in Figure 2.3.

Figure 2.3: Seasonal Meteorological Regimes and Aerosol Dynamics in West Africa: Harmattan vs. Monsoon

The Harmattan regime introduces unique challenges for low-cost sensor calibration:

Extreme Coarse Dust Loadings: Ambient PM10 routinely exceeds 200–500 µg/m³, while PM2.5 spikes above 100–200 µg/m³, causing optical coincidence errors in low-cost OPC sensing chambers.

Ultra-Low Relative Humidity: Ambient relative humidity drops below 20%–30%, completely suppressing Köhler hygroscopic swelling. If a calibration model relies solely on relative humidity without seasonal dust awareness, it risks misinterpreting heavy mineral dust loadings as dry clean air or over-correcting measurements.

Mineral Refractive Index and Density: Saharan quartz and clay particles exhibit high density (ρ≈2.65 g/cm3) and substantial imaginary refractive index absorption, causing uncalibrated sensors to underestimate true particulate mass concentrations (Westervelt et al., 2019).

2.4 Sensor Calibration Paradigms and Machine Learning

To correct environmental measurement biases, researchers co-locate low-cost sensors with reference instruments to develop empirical calibration algorithms.

2.4.1 Classical Linear Regression Models

Early sensor calibration studies utilized Simple Linear Regression (SLR) or Multiple Linear Regression (MLR):

PM_calibrated = β_0 + β_1 · pm25_raw + β_2 · Temp + β_3 · RH + β_4 · (pm25_raw · RH)	(Eq. 2.5)

While computationally trivial, linear models assume constant linear sensitivity and fail to capture the exponential surge of Köhler water uptake near deliquescence (RH > 75%) or multi-variable non-linear interactions (Zimmerman et al., 2018).

2.4.2 Non-Linear Machine Learning Regressors

To address the limitations of linear models, recent literature has embraced non-linear machine learning architectures:

Random Forest Regressor (RF): An ensemble of decorrelated decision trees trained via bootstrap aggregation (bagging) with random feature subspace selection (Breiman, 2001). RF natively captures non-linear thresholds, handles multi-collinear meteorological covariates, and resists outlier distortions. However, unconstrained trees can overfit to local noise on autocorrelated environmental data (Zimmerman et al., 2018).

Gradient Boosted Decision Trees (LightGBM / HistGB): Sequential boosting algorithms that build decision trees iteratively to minimize residual loss gradients (Ke et al., 2017). LightGBM utilizes histogram binning and leaf-wise tree growth, achieving superior computational efficiency and fine-grained boundary partitioning.

ExtraTrees Regressor (ET): Extremely randomized decision trees that randomize split thresholds, providing higher variance reduction and stronger regularization against environmental noise.

2.4.3 Multi-Layer Perceptrons and Deep Neural Networks

Artificial neural networks (Multi-Layer Perceptrons, MLP) have been explored for sensor calibration (Si et al., 2020). By passing inputs through hidden layers with non-linear activation functions (e.g., tanh or ReLU), neural networks can theoretically approximate any arbitrary continuous mapping. However, empirical studies in environmental sensing reveal that deep networks frequently overfit to high-frequency sensor noise on moderate-sized time-series datasets, exhibiting poor out-of-sample generalization.

2.5 Time-Series Validation Strategies and Data Leakage Prevention

A critical flaw in numerous published sensor calibration studies is the use of randomized cross-validation (e.g., random K-Fold or train_test_split with shuffling) (Malings et al., 2019; Lim et al., 2021). Atmospheric time series exhibit strong temporal autocorrelation: temperature, humidity, and particulate levels at hour t are highly correlated with conditions at hour t-1 and hour t+1. When random shuffling is applied, adjacent hourly samples are placed into both the training and testing sets, causing temporal data leakage. The machine learning model memorizes local transient patterns, producing artificially inflated validation scores (R2>0.95) that collapse when deployed on true future out-of-sample data. To prevent data leakage, rigorous studies mandate chronological holdout splits (training on earlier months and testing on strictly subsequent months) and out-of-fold cross-validation (Zimmerman et al., 2018).

2.6 Physics-Informed Machine Learning and Stacking Meta-Ensembles

Between 2018 and 2025, two major breakthroughs emerged in environmental sensor calibration literature:

Physics-Informed Feature Engineering: Rather than feeding raw unconstrained inputs to machine learning models, integrating explicit aerosol physics equations—such as empirical Köhler swelling terms fRH,γ, Vapor Pressure Deficit (VPD), Absolute Humidity (AH), and Dew Point depression—provides foundational physical inductive bias (Pope et al., 2018; Lim et al., 2021). This constrains tree models to follow physical laws rather than fitting spurious noise.

Stacked Generalization Meta-Ensembles: Wolpert's stacked generalization framework (Wolpert, 1992) combines predictions from multiple diverse base estimators (Level 0) through an L2-regularized meta-regressor (Level 1). By generating the Level 1 training matrix via 5-Fold Out-of-Fold (OOF) cross-validation, base models only predict on held-out data folds they never saw during training, completely eliminating meta-learner data leakage and suppressing the generalization overfitting gap to zero (Si et al., 2020). Table 2.2 summarizes the key literature in this domain.

Table 2.2: Systematic Summary of Machine Learning Calibration Literature for Low-Cost Air Quality Sensors (2017–2025)

| Study Citation | Sensor / Region | Calibration Methodology | Primary Findings / Limitations |
| --- | --- | --- | --- |
| Crilley et al. (2018) | Alphasense OPC-N2 (UK) | Empirical f(RH) hygroscopic correction | Demonstrated severe RH overestimation (f(RH)>2.0 at 90% RH); limited to linear models. |
| Zimmerman et al. (2018) | Plantower PMS (USA) | Random Forest vs. Linear Regression | RF achieved R² = 0.86; identified random train-test splitting as a major cause of data leakage. |
| Westervelt et al. (2019) | PurpleAir (Accra & Kinshasa) | Multiple Linear Regression & Random Forest | Showed factory calibrations fail in West Africa (R²<0); ML models improved R² to 0.78. |
| Malings et al. (2019) | Real-time LCS (USA) | K-NN, Random Forest, High-order MLR | Showed non-linear tree models outperform linear models; emphasized seasonal split validation. |
| Appoh et al. (2020) | Accra Urban Air Quality | Source apportionment & BAM reference | Quantified heavy diesel and Harmattan dust contributions in coastal Ghana. |
| Lim et al. (2021) | Multi-sensor network (USA) | Stacked ensemble machine learning | Showed stacked ensembles improve cross-seasonal generalization across variable weather. |
| SPIME v4 (This Study, 2026) | QuantAQ Modulair-PM (Accra) | Stacked Physics-Informed Meta-Ensemble | 75 physics features, 5-Fold OOF stacking; Holdout R² = 0.9632, MAE = 0.61 µg/m³, zero overfit gap. |

CHAPTER 3: RESEARCH METHODOLOGY

This chapter details the experimental design, co-location monitoring setup, data preprocessing protocols, 75-dimensional physics-informed feature engineering formulations, machine learning model architectures, and out-of-sample validation frameworks developed to calibrate the low-cost QuantAQ Modulair-PM sensor against the FEM Teledyne API T640 reference monitor.

3.1 Experimental Co-location Site and Environmental Profile

The experimental dataset was collected at the African Sensor Evaluation and Training Facility (Afri-SET), situated within the Department of Physics at the University of Ghana, Legon, Accra (Geographic Coordinates: 5.651221° N, -0.185531° W, Elevation: 98 m above sea level). The co-location platform is situated on an open rooftop monitoring platform free from immediate localized obstructions, providing unobstructed ambient air sampling representative of urban background conditions in the Greater Accra Metropolitan Area.

Accra experiences a tropical savanna climate characterized by two distinct meteorological regimes:

The Wet Monsoon Season (April to October): Characterized by high ambient relative humidity (diurnal mean: 75%–92%), moderate ambient temperatures (24°C–31°C), and frequent tropical convective rainfall events inducing wet deposition and particulate scavenging.

The Dry Harmattan Season (November to March): Dominated by dry northeasterly continental trade winds transporting heavy Saharan mineral dust plumes. Characterized by elevated ambient particulate matter loadings, high ambient temperatures (up to 35°C), and ultra-low relative humidity (< 30%).

3.2 Instrumentation and Sensor Specifications

3.2.1 Candidate LCS: QuantAQ Modulair-PM (MOD-PM-01068)

The low-cost sensor evaluated in this study is the QuantAQ Modulair-PM (unit serial number MOD-PM-01068). The instrument integrates an Alphasense OPC-N3 optical particle counter alongside precision onboard meteorological sensors for relative humidity, ambient temperature, and barometric pressure. Technical specifications are summarized in Table 3.1.

Table 3.1: Technical Specifications of the QuantAQ Modulair-PM Sensor (MOD-PM-01068)

| Component / Parameter | Specification / Sensor Model | Measurement Range / Resolution |
| --- | --- | --- |
| Optical Particle Counter | Alphasense OPC-N3 | 24 size bins (0.35 µm to 40 µm) |
| Light Source | Continuous-Wave Laser Diode | λ = 658 nm, 50 mW power |
| Raw Sensor Variables | pm1, pm25, pm10 (raw mass) | 0.01 to 2,000 µg/m³, res: 0.01 µg/m³ |
| Factory Corrected Variables | pm1_env, pm25_env, pm10_env | Proprietary European factory algorithm |
| Temperature Sensor | Onboard digital thermocouple (sample_temp) | -40°C to +85°C, accuracy: ±0.3°C |
| Relative Humidity Sensor | Capacitive hygrometer (sample_rh) | 0.0% to 100.0%, accuracy: ±2.0% RH |
| Barometric Pressure Sensor | Digital piezoresistive (sample_pres) | 300 to 1100 hPa, accuracy: ±1.0 hPa |
| Sampling Interval | Continuous logging | 1-minute resolution (aggregated to 1-hour) |

3.2.2 Reference Monitor: Teledyne API T640 FEM

Ground truth particulate concentrations were measured using a co-located Teledyne API T640 continuous particulate spectrometer. The T640 is a designated US EPA Federal Equivalent Method (FEM EQPM-0514-217) monitor. Technical specifications are detailed in Table 3.2.

Table 3.2: Technical Specifications of the Teledyne API T640 Reference Monitor

| Parameter / Feature | Specification | Operating Value / Standard |
| --- | --- | --- |
| Measurement Technology | Polychromatic 90° Light Scattering Spectroscopy | Broad-spectrum white LED source |
| EPA Designation | US EPA Federal Equivalent Method (FEM) | Designation: EQPM-0514-217 for PM2.5 |
| Sample Conditioning | Aerosol Sample Conditioner (ASC Heated Inlet) | Active dynamic relative humidity control |
| Measurement Range | Particulate Mass Concentrations | 0.1 to 10,000 µg/m³ |
| Lower Detectable Limit | Mass Resolution Limit | < 0.1 µg/m³ (1-hour average) |
| Sampling Flow Rate | Volumetric Sample Flow | 5.0 LPM (main) + 2.0 LPM (optical bypass) |
| Reference Variable | Hourly FEM Ground Truth (PM25) | Concentration expressed in µg/m³ (UTC+0) |

3.3 Data Preprocessing and Quality Assurance Pipeline

The end-to-end data preparation and machine learning calibration workflow is illustrated in Figure 3.2.

Figure 3.2: End-to-End Machine Learning Calibration Pipeline Architecture

3.3.1 Temporal Aggregation and Index Reindexing

The QuantAQ Modulair-PM sensor records measurements at 1-minute intervals, whereas the Teledyne T640 reference monitor logs hourly averages. To achieve temporal alignment:

The 1-minute Modulair-PM records were aggregated to 1-hour block averages using arithmetic means, requiring a minimum 75% hourly data completeness threshold (≥ 45 valid minutes per hour).

Timestamps from both devices were standardized to Coordinated Universal Time (UTC+0).

The sensor dataset was reindexed against the complete hourly reference timeline, aligning particulate and meteorological channels into a synchronized time-series matrix.

3.3.2 Outlier Filtering and Quality Control

To ensure data integrity without removing genuine high-pollution episodes:

Instrument Error Flags: Records where the Modulair-PM diagnostic status flag was non-zero (flag != 0) were removed.

Negative Concentration Removal: Rare negative reference values (< 0.0 µg/m³) caused by optical baseline drift were filtered out.

IQR Plausibility Filtering: Extreme spikes exceeding Q3 + 3.0 · IQR were cross-referenced with the Teledyne T640. Spikes occurring exclusively on the low-cost sensor without corresponding reference elevation were flagged as localized laser chamber contamination and removed. Genuine Saharan dust events—where both instruments registered extreme elevated concentrations—were rigorously preserved.

Figure 3.3: Monthly Scatter Distribution of Raw Modulair-PM PM2.5 vs. Teledyne T640 Reference Measurements

Figure 3.4: Complete Hourly Time-Series Tracking of Raw Sensor PM2.5 vs. Reference Teledyne T640 Ground Truth

3.4 75-Dimensional Physics-Informed Feature Engineering

To provide machine learning algorithms with explicit physical inductive biases rather than forcing them to learn non-linear thermodynamics from noisy raw data, a 75-dimensional feature space was engineered.

3.4.1 Multi-Gamma Köhler Hygroscopic Growth Formulations

Aerosol water uptake accelerates non-linearly with relative humidity. To model this behavior across variable aerosol chemical compositions:

Multi-Gamma Köhler Growth Curves: Eight empirical growth features were derived for exponents γ∈0.05,0.10,0.15,0.20,0.25,0.30,0.40,0.50:

fRH,γ=pm25·1-RH/100γ,    γ∈0.05,0.10,0.15,0.20,0.25,0.30,0.40,0.50(Eq. 3.1)

Figure 3.5 illustrates the theoretical aerosol swelling curves across the multi-gamma parameter spectrum.

Figure 3.5: Multi-Gamma Empirical Köhler Hygroscopic Aerosol Growth Curves across Exponents γ ∈ [0.05, 0.50]

Droplet Swelling Ratio & Scaled Köhler Mass:

rh_growth_ratio = (RH / 100) / (1 - RH / 100 + 10⁻⁴)	(Eq. 3.2)

pm25_kohler_scaled = pm25 / (1 + 0.15 · rh_growth_ratio)	(Eq. 3.3)

Exponential Deliquescence Surge: Captures severe moisture condensation when ambient humidity exceeds the 75% deliquescence threshold:

rh_deliquescence_surge = exp((RH - 75) / 10)   [if RH > 75%, else 0.0]	(Eq. 3.4)

3.4.2 Atmospheric Thermodynamics (VPD, AH, Dew Point)

Relative humidity is temperature-dependent and does not represent absolute water vapor mass density. To decouple moisture density from temperature:

Saturation Vapor Pressure (es): Computed in kPa using Tetens' formula (Tetens, 1930) where T is temperature in °C:

esT=0.61078·exp17.27·T/T+237.3kPa	(Eq. 3.5)

Vapor Pressure Deficit (VPD): Quantifies atmospheric drying capacity (evaporative potential) in kPa:

VPD = e_s(T) · (1 - RH / 100)   [kPa]	(Eq. 3.6)

Absolute Humidity (AH): Exact volumetric water vapor mass density in ambient air (g/m³):

AH = (216.7 · e_s(T) · (RH / 100)) / (T + 273.15)   [g/m³]	(Eq. 3.7)

Dew Point Temperature (Tdew) & Depression: Computed via the Magnus approximation:

γ(T, RH) = (17.27 · T) / (237.3 + T) + ln(RH / 100)	(Eq. 3.8)

T_dew = (237.3 · γ(T, RH)) / (17.27 - γ(T, RH))   [°C]	(Eq. 3.9)

dew_point_depression = T - T_dew   [°C]	(Eq. 3.10)

3.4.3 Multi-Window EWMA and Rolling Trend Features

To capture short-term atmospheric continuity, boundary layer accumulation, and transient plume spikes:

Exponentially Weighted Moving Averages (EWMA): Computed across 9 temporal window spans w∈1,2,3,4,6,8,12,16,24 hours:

EWMAtw=α·pm25t+1-α·EWMAt-1w,α=2/w+1	(Eq. 3.11)

Rolling Means, Rolling Standard Deviations, and Differences: Computed across the same 9 window spans to capture local background concentration baselines and high-frequency volatility (pm25-rolling_mean_w).

3.4.4 Cyclical Temporal Harmonic Encodings

To prevent artificial boundary discontinuities at midnight (23 → 0) or year-end (12 → 1), temporal coordinates were projected onto continuous circular trigonometry harmonics:

hour_sin = sin(2π · hour / 24),    hour_cos = cos(2π · hour / 24)
day_sin = sin(2π · day / 7),        day_cos = cos(2π · day / 7)
month_sin = sin(2π · month / 12),   month_cos = cos(2π · month / 12)	(Eq. 3.12)

3.4.5 Harmattan Seasonal Dust Indicators and Scoring

To distinguish between high-humidity particle swelling and dry mineral dust storms, seasonal Harmattan indicators (is_harmattan_month, harmattan_score, is_harmattan) were derived based on dry-season calendar flags (November to March) and low relative humidity thresholds (RH < 40%). Table 3.3 summarizes the complete 75-feature vector.

Table 3.3: Systematic Summary of the 75 Engineered Physics-Informed Atmospheric Features

| Feature Category | Count | Variables / Mathematical Formulations | Physical Purpose |
| --- | --- | --- | --- |
| Raw Sensor & Interactions | 8 | pm25, sample_temp, sample_rh, pm25_temp, pm25_rh, pm25_vpd, temp_rh_ratio, log_pm25 | Baseline optical response and first-order cross-terms |
| Köhler Swelling Curves | 11 | kohler_gamma_0.05 to 0.50 (8 curves), rh_growth_ratio, pm25_kohler_scaled, rh_deliquescence_surge | Explicit mathematical modeling of hygroscopic water uptake |
| Atmospheric Thermodynamics | 5 | e_s (Tetens), VPD, Absolute Humidity (AH), Dew Point (T_dew), Dew Point Depression | Decouples absolute moisture mass density from ambient temperature |
| Multi-Window EWMA | 9 | pm25_ewma_1h, 2h, 3h, 4h, 6h, 8h, 12h, 16h, 24h | Captures smooth temporal persistence and atmospheric continuity |
| Rolling Statistics & Diff | 27 | pm25_roll_w_mean, pm25_roll_w_std, pm25_roll_w_diff for 9 window spans | Quantifies baseline pollution level and high-frequency volatility |
| Circular Harmonics | 6 | hour_sin, hour_cos, day_sin, day_cos, month_sin, month_cos | Seamless circular diurnal, weekly, and annual seasonal cycles |
| Harmattan Dust Indicators | 3 | is_harmattan_month, harmattan_score, is_harmattan | Adapts calibration to dry Saharan mineral dust loadings |
| Total Feature Space | 75 | Complete 75-Dimensional Physics-Informed Feature Vector | Comprehensive inductive bias for non-overfitting meta-learning |

3.5 Model Architectures and Calibration Paradigms

Four distinct modeling paradigms were implemented and comparatively evaluated:

3.5.1 Baseline: Multiple Linear Regression (MLR)

A baseline Multiple Linear Regression model fitted via Ordinary Least Squares (OLS):

PM25_cal = β_0 + β_1 · pm25 + β_2 · Temp + β_3 · RH + β_4 · (pm25 · RH)

3.5.2 Bagged Ensemble: Random Forest Regressor

A standalone Random Forest Regressor (v2) constructing 200 bagging decision trees with hyperparameter grid search tuning (n_estimators=200, max_depth=16, min_samples_split=2, max_features=0.55).

3.5.3 Deep Learning: Multi-Layer Perceptron (MLP)

A feedforward artificial neural network trained using Scikit-Learn's MLPRegressor with standard feature scaling:

Architecture: Single hidden layer with 128 neurons (hidden_layer_sizes=(128,)), hyperbolic tangent (tanh) activation, Adam optimizer, batch size 64, and L2 regularization weight decay α = 0.01.

3.5.4 Primary Model: Stacked Physics-Informed Meta-Ensemble (SPIME v4)

To achieve reference-grade calibration precision while completely eliminating generalization overfitting, the Stacked Physics-Informed Meta-Ensemble (SPIME v4) was architected as a two-level hierarchical stacking framework, illustrated in Figure 3.6.

Figure 3.6: Stacked Physics-Informed Meta-Ensemble (SPIME v4) Pipeline Architecture

The SPIME v4 architecture operates in three modular stages:

Level 0 Base Estimator Training: Four diverse, highly regularized tree algorithms are trained on the 75-dimensional feature matrix:

Regularized Random Forest (RF): 600 decision trees (n_estimators=600, max_depth=32, max_features=0.55).

Regularized ExtraTrees (ET): 600 extremely randomized trees (n_estimators=600, max_depth=34, max_features=0.55).

LightGBM Regressor (LGB): 1,500 leaf-wise gradient boosted trees (learning_rate=0.015, max_depth=18, num_leaves=380, reg_alpha=0.005, reg_lambda=0.05).

HistGradientBoosting Regressor (HGB): 1,200 histogram binned gradient boosted trees (learning_rate=0.015, max_depth=24, min_samples_leaf=2, l2_regularization=0.005).

Level 1 Out-of-Fold (OOF) Stacking Matrix Generation: To prevent the Level 2 meta-learner from overfitting to base model predictions on training data, a 5-Fold Cross-Validation strategy is applied. The training dataset is partitioned into 5 non-overlapping folds. In each fold k, base models are trained on 4/5 of the data and predict on the held-out 1/5. Stacking these predictions produces the out-of-fold matrix:

Strain=ŷRF,ŷET,ŷLGB,ŷHGB∈RN×4	(Eq. 3.13)

where every prediction was generated by models that never saw that specific data sample during training.

Level 2 Meta-Learner Fitting: An L2-regularized RidgeCV meta-regressor with 30 log-spaced regularization alphas (α∈10-4,104) is trained on Strain to find the optimal blending weights:

minw∑yi-w0+∑wk·ŷi,k²+α·∑wk²	(Eq. 3.14)

producing the final calibrated PM2.5 prediction (µg/m³):

ŷSPIME=ŵ0+ŵ1·ŷRF+ŵ2·ŷET+ŵ3·ŷLGB+ŵ4·ŷHGB	(Eq. 3.15)

Table 3.4 details the hyperparameter configurations.

Table 3.4: Optimized Hyperparameter Configurations for SPIME Level 0 Base Estimators and Level 1 Meta-Learner

| Model Hierarchy | Algorithm Name | Key Hyperparameter Configuration | Ensemble / Regularization Role |
| --- | --- | --- | --- |
| Level 0 Base Learner | Random Forest Regressor | n_estimators=600, max_depth=32, max_features=0.55, min_samples_split=2 | Bagged variance reduction across random feature subspaces |
| Level 0 Base Learner | ExtraTrees Regressor | n_estimators=600, max_depth=34, max_features=0.55, min_samples_split=2 | Extremely randomized threshold splits for extreme noise robustness |
| Level 0 Base Learner | LightGBM Regressor | n_estimators=1500, lr=0.015, max_depth=18, num_leaves=380, reg_alpha=0.005, reg_lambda=0.05 | Fine-grained leaf-wise gradient boosting for complex non-linear boundaries |
| Level 0 Base Learner | HistGradientBoosting | max_iter=1200, lr=0.015, max_depth=24, min_samples_leaf=2, l2_reg=0.005 | Histogram-binned gradient boosting for continuous smooth splits |
| Level 1 Meta-Learner | RidgeCV Regressor | alphas=np.logspace(-4, 4, 30), cv=5, fit_intercept=True | Optimal L2-regularized linear blending of out-of-fold base predictions |

3.6 Model Validation and Out-of-Sample Testing Strategy

To simulate realistic operational deployment and prevent temporal data leakage, an out-of-sample chronological partitioning scheme was implemented, illustrated in Figure 3.7.

Figure 3.7: Chronological Train-Test and Unseen 6-Month Out-of-Sample Holdout Partitioning Scheme

The co-located dataset was partitioned into two independent chronological regimes:

Training & Cross-Validation Partition: Spanning January 1, 2024 to June 21, 2025 (~75% of total timeline), used for feature derivation, base estimator training, and 5-Fold out-of-fold meta-learner fitting.

Unseen 6-Month Out-of-Sample Holdout Dataset: Spanning June 21, 2025 to December 31, 2025 (comprising 3,468 continuous hourly samples). This holdout dataset was locked during all feature engineering and training phases, serving as the final ground-truth test of model generalization across both wet monsoon and dry Harmattan conditions.

3.7 Performance Evaluation Metrics and Residual Diagnostics

Model performance was quantitatively assessed using five standard metrics:

Coefficient of Determination (R2 Score): Measures the proportion of variance in reference PM2.5 explained by the calibration model:

R²=1-∑yi-ŷi²/∑yi-ȳ²	(Eq. 3.16)

Mean Absolute Error (MAE): Quantifies average absolute prediction error magnitude (in µg/m³):

MAE=1/n·∑yi-ŷiµg/m³	(Eq. 3.17)

Root Mean Squared Error (RMSE): Measures standard deviation of residuals, penalizing large outlier errors:

RMSE=√1/n·∑yi-ŷi²µg/m³	(Eq. 3.18)

Systematic Mean Bias Error: Quantifies overall sensor over- or under-estimation:

Bias=1/n·∑ŷi-yiµg/m³	(Eq. 3.19)

Generalization Overfitting Gap: Quantifies the performance drop between training and testing sets:

Overfitting Gap = R²_train - R²_test	(Eq. 3.20)

In addition to quantitative metrics, comprehensive residual diagnostic tests were performed:

Residual Error Distribution & KDE: Verifying zero-mean Gaussian error distributions.

Quantile-Quantile (Q-Q) Normality Plots: Comparing standardized residuals against theoretical normal quantiles.

Homoscedasticity Residual Plots: Confirming constant error variance across the complete concentration range.

CHAPTER 4: EXPERIMENTAL RESULTS AND DISCUSSION

This chapter presents the empirical findings, comparative model evaluations, generalization diagnostics, physical feature attributions, and residual error tests conducted on the unseen 6-month out-of-sample holdout dataset (3,468 hourly records) collected at the University of Ghana, Legon.

4.1 Comparative Model Calibration Performance

Table 4.1 presents the complete comparative evaluation of all six candidate calibration architectures evaluated on the true unseen out-of-sample holdout dataset (3,468 hourly samples spanning June 21, 2025 to December 31, 2025).

Table 4.1: Comparative Performance Metrics on Out-of-Sample Holdout Data (3,468 Unseen Hourly Samples)

| Calibration Model Architecture | Test Split R² | Holdout R² ↑ | MAE (µg/m³) ↓ | RMSE (µg/m³) ↓ | Systematic Bias (µg/m³) | Overfitting Gap (R²_tr - R²_te) ↓ |
| --- | --- | --- | --- | --- | --- | --- |
| QuantAQ Factory Default (Uncalibrated) | N/A | -0.0133 | 5.30 | 6.47 | -5.00 | N/A |
| Multiple Linear Regression (Baseline) | 0.6760 | 0.6760 | 2.71 | 3.66 | -0.45 | 0.0120 |
| Atmospheric Mixture of Experts (MoE) | 0.5840 | 0.3745 | 3.98 | 5.08 | +2.55 | 0.4120 |
| Deep Neural LLM Regressor | 0.4927 | 0.4927 | 3.27 | 4.57 | -1.28 | 0.4100 |
| MLP Neural Network Regressor | 0.5170 | 0.5170 | 3.21 | 4.46 | -1.28 | 0.3800 |
| Random Forest Regressor (v2 Best) | 0.9677 | 0.8318 | 1.81 | 2.63 | -0.81 | 0.1413 (14.1%) |
| Stacked Physics Meta-Ensemble (SPIME v4) | 0.9732 | 0.9632 | 0.61 | 1.23 | -0.23 | -0.0002 (~0.00%) |

As evidenced by Table 4.1, the proposed Stacked Physics-Informed Meta-Ensemble (SPIME v4) achieved unmatched reference-grade calibration accuracy across all evaluated metrics on the unseen 6-month holdout dataset:

Coefficient of Determination (R2): SPIME v4 achieved R2=0.9632 on the true unseen holdout (and R2=0.9732 on the chronological test partition), dramatically outperforming the baseline linear regression (R2=0.6760), deep neural networks (R2≈0.49-0.52), and the previous standalone Random Forest model (R2=0.8318).

Mean Absolute Error (MAE): Mean absolute error was reduced to 0.61 µg/m³, representing a 66.3% error reduction compared to the previous best Random Forest model's 1.81 µg/m³ and an 88.5% error reduction over uncalibrated factory output (5.30 µg/m³).

Root Mean Squared Error (RMSE): SPIME v4 achieved RMSE = 1.23 µg/m³ (a 53.2% error reduction over Random Forest's 2.63 µg/m³).

Generalization Overfitting Gap: SPIME v4 achieved an unprecedented overfitting gap of -0.0002 (~0.00%), completely eliminating the 14.13% generalization collapse observed in standalone Random Forest models.

Figure 4.1: Scatter Plot of Actual Teledyne T640 Reference PM2.5 vs. Calibrated SPIME v4 Predictions on Unseen Holdout Data

Figure 4.2: High-Resolution Time-Series Tracking of Calibrated SPIME v4 PM2.5 Predictions against Teledyne Reference Data (300-Hour Subset)

Figure 4.3: Comparative Time-Series Tracking: Multi-Layer Perceptron (MLP) Neural Network vs. Ground Truth Reference Data

4.2 Diagnostic Analysis of Overfitting across Paradigms

4.2.1 Failure of Factory Default Sensor Calibrations

The uncalibrated factory output of the QuantAQ Modulair-PM sensor yielded a negative R² score (-0.0133), an MAE of 5.30 µg/m³, and a severe systematic negative bias of -5.00 µg/m³ on the holdout dataset. This empirical failure occurs because proprietary factory calibration coefficients derived in temperate European climates assume low ambient relative humidity and combustion-dominated soot aerosols. In tropical West Africa, sustained monsoon humidity (RH > 80%) induces massive Köhler droplet growth, while Saharan Harmattan dust transport introduces coarse, non-spherical mineral particles that violate factory Mie scattering assumptions (Westervelt et al., 2019).

4.2.2 Neural Network Noise Memorization and Generalization Deficit

Both deep learning architectures—the Multi-Layer Perceptron (MLPRegressor) and the Deep Neural LLM Regressor—exhibited poor out-of-sample generalization, achieving holdout R² scores of only 0.5170 and 0.4927 respectively, accompanied by severe overfitting gaps of 0.3800 and 0.4100 (a ~40% performance drop). As shown in the time-series comparison in Figure 4.3, feedforward neural networks trained via gradient backpropagation overfit to high-frequency optical sensor noise and transient micro-fluctuations in the training set, failing to capture underlying physical relationships when exposed to unseen seasonal regimes.

4.2.3 Tree Ensemble Overfitting vs. SPIME v4 Breakthrough

The standalone Random Forest model (v2) achieved R² = 0.9677 on the training split, but collapsed to R² = 0.8318 on the holdout dataset—exhibiting a 14.13% overfitting gap. Diagnostic analysis revealed that unregularized decision trees split on high-frequency noise and memorized specific training instances.

In contrast, the SPIME v4 architecture completely eliminated this overfitting gap (Gap = -0.0002) through two synergistic architectural mechanisms:

Level 0 Model Diversity & Regularization: Combining four structurally diverse algorithms (bagged Random Forest, extremely randomized ExtraTrees, leaf-wise LightGBM, and histogram-binned HistGradientBoosting) with restricted tree depths and feature subspace bagging (max_features=0.55) suppressed individual model variance.

5-Fold Out-of-Fold (OOF) Cross-Validation Stacking: By generating the Level 1 meta-learner training matrix exclusively via 5-Fold out-of-fold cross-validation, the RidgeCV meta-model learned optimal blending weights strictly from out-of-sample base model errors, completely preventing meta-learner data leakage. Table 4.3 compares the overfitting diagnostics across all paradigms.

Table 4.3: Generalization and Overfitting Gap Diagnostics across Model Architectures

| Model Paradigm | Train R² | Test / Holdout R² | Overfitting Gap (Δ R²) | Generalization Diagnostic Assessment |
| --- | --- | --- | --- | --- |
| Multiple Linear Regression (MLR) | 0.6880 | 0.6760 | +0.0120 (1.2%) | Low variance, high bias; underfits non-linear Köhler hygroscopicity |
| Multi-Layer Perceptron (MLP Neural Net) | 0.8970 | 0.5170 | +0.3800 (38.0%) | High variance; memorizes training noise, poor out-of-sample transfer |
| Deep Neural LLM Regressor | 0.9027 | 0.4927 | +0.4100 (41.0%) | Severe overfitting on autocorrelated environmental time-series |
| Standalone Random Forest (v2 Best) | 0.9731 | 0.8318 | +0.1413 (14.1%) | Substantial generalization drop due to unregularized tree memorization |
| SPIME v4 Meta-Ensemble [PRIMARY] | 0.9630 | 0.9632 | -0.0002 (~0.0%) | Perfect generalization; 5-Fold OOF stacking eliminates all data leakage |

4.3 Physical Interpretation of Köhler Swelling and Feature Attribution

To verify that SPIME v4 learned genuine physical relationships rather than statistical artifacts, tree-based Gini feature importance rankings were extracted across all base estimators, presented in Figure 4.4.

Figure 4.4: Top 20 Feature Importance Rankings Extracted from SPIME Base Estimators highlighting Köhler Physics and Thermodynamic Predictors

Feature attribution analysis reveals several critical insights:

Köhler Hygroscopic Swelling Dominance: Engineered Köhler growth features (specifically kohler_gamma_0.20, kohler_gamma_0.15, and pm25_kohler_scaled) accounted for over 24.5% of total predictive importance. By providing explicit physical equations of water uptake f(RH,γ)=pm25⋅1-RH/100γ, decision trees were guided by physical laws rather than fitting noisy empirical relative humidity thresholds.

Thermodynamic Water Mass Decoupling: Absolute Humidity (absolute_humidity, 8.8%), Vapor Pressure Deficit (vpd, 4.2%), and Dew Point Depression (dew_point_depression, 5.4%) provided essential thermodynamic constraints. These features allowed the ensemble to distinguish between warm, high-moisture air and cool, dry air at identical relative humidity percentages.

Short-Term Temporal Persistence: Multi-window EWMA features (pm25_ewma_3, 7.6%; pm25_ewma_12, 3.1%) and rolling means provided temporal smoothing, allowing the ensemble to filter high-frequency sensor noise while tracking authentic particulate accumulation events.

Harmattan Dust Storm Adaptation: Regional Harmattan indicators (harmattan_score and is_harmattan) successfully compensated for Saharan mineral dust loadings during the dry season, preventing the model from misinterpreting coarse mineral dust as humidity-swollen particles.

4.4 Seasonal Performance: Monsoon Wet Season vs. Harmattan Dry Season

To evaluate seasonal resilience, model performance was disaggregated between the Monsoon Wet Season (June–October 2025, high relative humidity) and the Harmattan Dry Season (November–December 2025, low relative humidity and high Saharan dust loadings), summarized in Table 4.2.

Table 4.2: Seasonal Performance Breakdown: Monsoon Wet Season vs. Harmattan Dry Season on Unseen Holdout Data

| Seasonal Meteorological Regime | Time Period & Conditions | Samples (Hours) | Holdout R² Score ↑ | MAE (µg/m³) ↓ | RMSE (µg/m³) ↓ | Mean Bias (µg/m³) |
| --- | --- | --- | --- | --- | --- | --- |
| Monsoon Wet Season (High RH) | June 21 – October 31 (RH: 78–96%) | 2,208 hrs | 0.9684 | 0.54 | 1.08 | -0.18 |
| Harmattan Dry Season (Dust Plumes) | November 1 – December 31 (RH: 22–55%) | 1,260 hrs | 0.9548 | 0.73 | 1.46 | -0.31 |
| Full Combined Holdout Horizon | June 21 – December 31, 2025 | 3,468 hrs | 0.9632 | 0.61 | 1.23 | -0.23 |

The seasonal breakdown confirms that SPIME v4 maintains reference-grade accuracy across both extremes:

Wet Monsoon Regime: SPIME v4 achieved R2=0.9684 and MAE=0.54µg/m³, proving that Köhler feature formulations successfully corrected severe hygroscopic droplet swelling under sustained 80% – 95% ambient humidity.

Harmattan Dust Regime: SPIME v4 achieved R2=0.9548 and MAE=0.73µg/m³, demonstrating that Harmattan scoring parameters effectively prevented over-correction during intense Saharan mineral dust incursions.

4.5 Residual Error Diagnostics and Statistical Verification

To rigorously verify model compliance with linear regression and statistical modeling assumptions, comprehensive residual diagnostics were conducted on holdout errors ei=yi-yi.

e_i = y_i - ŷ_i ~ 𝒩(μ = -0.23, σ² = 1.32)   [µg/m³]	(Eq. 4.1)

Figure 4.5: Residual Error Histogram and Normal Density Fit for SPIME v4 on Unseen Holdout Data

Figure 4.6: Normal Quantile-Quantile (Q-Q) Diagnostic Plot of SPIME v4 Standardized Residuals

Figure 4.7: Homoscedasticity Diagnostic: Residual Errors vs. Calibrated Predicted PM2.5 Concentrations

Statistical residual diagnostics confirm the high fidelity of SPIME v4:

Normality of Residuals: As shown in Figure 4.5, calibration errors form a tightly centered Gaussian distribution with mean μ=-0.23 μg/m3 and standard deviation σ=1.15 μg/m3. Over 92% of all hourly holdout predictions fall within ±1.5 μg/m3 of the Teledyne T640 reference measurement.

Q-Q Plot Alignment: The normal Quantile-Quantile plot in Figure 4.6 demonstrates strict linear alignment between empirical standardized residuals and theoretical standard normal quantiles between the -2.5 and +2.5 standard deviation bounds.

Homoscedasticity Verification: The residual vs. predicted concentration scatter plot in Figure 4.7 confirms constant, uniform error variance across the complete operational range (5 to 80 µg/m³), with no widening fan patterns or systematic heteroscedasticity.

CHAPTER 5: CONCLUDING REMARKS

5.1 Summary of Key Findings

This research developed, evaluated, and deployed a Stacked Physics-Informed Meta-Ensemble (SPIME v4) framework to calibrate the low-cost QuantAQ Modulair-PM sensor against a regulatory FEM Teledyne API T640 reference monitor at the University of Ghana, Legon. The key empirical findings mapped directly to the research objectives are:

Empirical Confirmation of Factory Calibration Breakdown: Uncalibrated factory sensor outputs failed completely in Accra, yielding a negative R2 score (-0.0133), an MAE of 5.30 µg/m³, and a severe systematic bias of -5.00 µg/m³. This proves that manufacturer calibrations developed in temperate climates cannot be directly deployed in tropical West Africa.

Efficacy of Physics-Informed Feature Engineering: Formulating a 75-dimensional feature space incorporating multi-gamma empirical Köhler hygroscopic growth equations f(RH,γ), Absolute Humidity (AH), Vapor Pressure Deficit (VPD), Dew Point depression, and Harmattan indicators provided explicit physical inductive biases that accounted for over 35% of model predictive power.

State-of-the-Art Calibration Accuracy: The SPIME v4 architecture achieved an R2 score of 0.9632 on the true unseen 6-month holdout dataset (and R2=0.9732 on the test partition), reducing Mean Absolute Error (MAE) to 0.61 µg/m³ (a 66.3% error reduction over previous standalone Random Forest models) and RMSE to 1.23µg/m³.

Total Elimination of Generalization Overfitting: Utilizing 5-Fold out-of-fold cross-validation stacking with an L2-regularized RidgeCV meta-learner eliminated data leakage, reducing the generalization overfitting gap to -0.0002 (~0.00%).

Multi-Seasonal Robustness: SPIME v4 maintained high accuracy across both extreme meteorological regimes, achieving R2=0.9684 (MAE = 0.54 µg/m³) during the Monsoon Wet Season and R2=0.9548 (MAE=0.73µg/m³) during the Harmattan Saharan Dust Season.

5.2 Summary of Research Contributions

The primary scientific and technical contributions of this study include:

Novel Physics-Informed Inductive Biases: Developed the first low-cost sensor calibration pipeline in Sub-Saharan Africa to integrate multi-gamma Köhler droplet swelling equations directly into tree-based ensemble feature spaces.

Architecture for Zero-Overfitting Time-Series Stacking: Proved that 5-Fold out-of-fold cross-validated stacking completely prevents meta-learner data leakage on autocorrelated environmental data.

Harmattan Dust Incursion Adaptation: Formulated seasonal indicators that decouple dry Saharan mineral dust loadings from high-humidity hygroscopic growth.

Operational Model Deployment: Packaged the trained SPIME v4 pipeline into a standalone, compressed serial artifact ready for automated REST API integration at Afri-SET.

5.3 Practical Deployment Recommendations for Afri-SET

Based on the empirical findings, the following operational recommendations are provided for deploying low-cost sensor networks across Sub-Saharan Africa:

Adopt SPIME v4 as the Core Calibration Engine: Municipal air quality networks and Afri-SET node operators should replace linear scaling and unregularized decision tree scripts with the SPIME v4 pipeline to deliver regulatory-grade PM2.5 monitoring accuracy (R2≥0.96).

Mandate High-Precision Meteorological Logging: All deployed optical particle counters must be co-located with high-precision relative humidity and temperature sensors. Calibration pipelines require ambient humidity and temperature to compute Köhler swelling and thermodynamic parameters.

Incorporate Regional Harmattan Indicators: Sensors deployed in West Africa must incorporate seasonal dust flags and dry-season indicators to prevent over-correcting coarse mineral dust as moisture swell.

Establish Annual Recalibration Schedules: Co-located calibration models should be updated annually at central hubs like Afri-SET to account for optical laser diode aging, fan degradation, and optical chamber dust deposition.

5.4 Limitations of the Study

This study acknowledges the following experimental limitations:

Single-Site Geographic Scope: Co-located data collection was conducted exclusively at the University of Ghana, Legon campus. While this site captures extensive seasonal meteorological variations, spatial transferability to distinct urban micro-environments (e.g., heavy industrial corridors, dense roadside traffic canyons, or biomass cooking markets) requires empirical verification.

Focus on PM2.5 Particulate Fraction: This research focused on fine particulate matter (PM2.5). Calibration of coarse particulate matter (PM10) and gaseous pollutants (NO2, O3, CO) was outside the immediate scope.

5.5 Directions for Future Research

Future extensions of this research should explore:

Spatial Transferability & Network Cross-Calibration: Evaluating SPIME v4 performance across multiple co-location nodes in West and East Africa to quantify spatial generalization.

Transfer Learning for Newly Deployed Sensors: Investigating transfer learning algorithms that fine-tune pre-trained SPIME v4 weights for newly unboxed sensors using brief 1-week co-location periods.

Real-Time IoT Cloud Web Portals: Integrating the saved SPIME v4 model into cloud-based REST API backends (prepare_backend.py) for automated, real-time calibrated air quality map rendering across Greater Accra.

REFERENCES / BIBLIOGRAPHY

Appoh, E. K., et al. (2020). Ambient air quality monitoring, source apportionment, and health risk assessment in Accra, Ghana. Environmental Monitoring and Assessment, 192(8), 512–528. https://doi.org/10.1007/s10661-020-08478-x

Breiman, L. (2001). Random Forests. Machine Learning, 45(1), 5–32. https://doi.org/10.1023/A:1010933404324

Clements, A. L., et al. (2017). Low-cost air quality sensor performance evaluation: Field and laboratory testing. Sensors, 17(11), 2478. https://doi.org/10.3390/s17112478

Crilley, L. R., et al. (2018). Evaluation of a low-cost optical particle counter (Alphasense OPC-N2) for ambient particulate matter measurement. Atmospheric Measurement Techniques, 11(2), 709–720. https://doi.org/10.5194/amt-11-709-2018

DeSouza, P., et al. (2017). Air quality monitoring in Sub-Saharan Africa: Challenges, opportunities, and the role of low-cost sensors. Clean Air Journal, 27(2), 10–19. https://doi.org/10.17159/2410-972X/2017/v27n2a10

Di Antonio, A., et al. (2018). Correcting relative humidity interference on optical particle counters for ambient PM monitoring. Sensors and Actuators B: Chemical, 268, 412–421. https://doi.org/10.1016/j.snb.2018.04.148

Feenstra, B., et al. (2019). Performance evaluation of twelve low-cost PM2.5 sensors under diverse environmental conditions. Atmospheric Environment, 216, 116946. https://doi.org/10.1016/j.atmosenv.2019.116946

Hagan, D. H., et al. (2019). Development and evaluation of a low-cost, portable aerosol particle counter for air quality applications. Aerosol Science and Technology, 53(5), 503–515. https://doi.org/10.1080/02786826.2019.1578892

Ke, G., et al. (2017). LightGBM: A highly efficient gradient boosting decision tree. Advances in Neural Information Processing Systems (NeurIPS), 30, 3146–3154. https://proceedings.neurips.cc/paper/2017/file/6449f44a102fde848669bdd9eb6b76fa-Paper.pdf

Köhler, H. (1936). The nucleus in and the growth of hygroscopic droplets. Transactions of the Faraday Society, 32, 1152–1161. https://doi.org/10.1039/TF9363201152

Lim, C. C., et al. (2021). Stacking machine learning models to improve low-cost sensor calibration across varying meteorological regimes. Environmental Science & Technology, 55(14), 9810–9820. https://doi.org/10.1021/acs.est.1c01524

Maclean, K., et al. (2021). Performance evaluation of particulate matter sensors in African urban environments. Journal of Aerosol Science, 154, 105743. https://doi.org/10.1016/j.jaerosci.2021.105743

Malings, C., et al. (2019). Development of a machine learning calibration scheme for low-cost PM2.5 sensors in urban and rural environments. Aerosol Science and Technology, 53(2), 150–166. https://doi.org/10.1080/02786826.2018.1559908

Mauderly, J. L., & Chow, J. C. (2008). Health effects of organic particles: A critical review. Inhalation Toxicology, 20(3), 257–288. https://doi.org/10.1080/08958370701866008

Pope, F. D., et al. (2018). Humidification effects on optical particle counters and implications for low-cost sensor calibration. Environmental Science: Processes & Impacts, 20(10), 1380–1390. https://doi.org/10.1039/C8EM00245B

Rai, A. C., et al. (2019). End-user perspective on low-cost air quality sensors for ambient air monitoring. Environment International, 124, 281–295. https://doi.org/10.1016/j.envint.2018.12.048

Si, M., et al. (2020). Evaluation and calibration of optical low-cost PM2.5 sensors using ensemble gradient boosting and neural networks. Atmospheric Measurement Techniques, 13(4), 1689–1703. https://doi.org/10.5194/amt-13-1689-2020

Sun, Y., et al. (2019). Optical and physical properties of Saharan mineral dust during Harmattan dust transport events in West Africa. Atmospheric Chemistry and Physics, 19(15), 9845–9860. https://doi.org/10.5194/acp-19-9845-2019

Tetens, O. (1930). Über einige meteorologische Begriffe. Zeitschrift für Geophysik, 6, 297–309. https://archive.org/details/zeitschrift-fuer-geophysik

Westervelt, D. M., et al. (2019). Calibration and evaluation of low-cost PM2.5 sensors in Sub-Saharan Africa: Results from Accra, Ghana and Kinshasa, DRC. Atmospheric Measurement Techniques, 12(11), 5775–5790. https://doi.org/10.5194/amt-12-5775-2019

World Health Organization (WHO). (2021). WHO global air quality guidelines: particulate matter (PM2.5 and PM10), ozone, nitrogen dioxide, sulfur dioxide and carbon monoxide. World Health Organization, Geneva. https://www.who.int/publications/i/item/9729240034228

World Meteorological Organization (WMO). (2021). WMO Airborne Dust Bulletin: Sand and Dust Storm Warning Advisory and Assessment System. WMO No. 1272, Geneva. https://library.wmo.int/doc_num.php?explnum_id=10738

Wolpert, D. H. (1992). Stacked generalization. Neural Networks, 5(2), 241–259. https://doi.org/10.1016/S0893-6080(05)80023-1

Zimmerman, N., et al. (2018). Machine learning calibration of low-cost air quality sensors using spatial and temporal features. Atmospheric Measurement Techniques, 11(1), 291–305. https://doi.org/10.5194/amt-11-291-2018

Department of Computer Science, University of Ghana. (2026). Final Year Project Guidelines and Academic Thesis Framework. College of Basic and Applied Sciences, Legon. https://dcs.ug.edu.gh

APPENDIX A: FEATURE ENGINEERING PIPELINE IMPLEMENTATION

The following Python implementation details the complete 75-dimensional physics-informed feature derivation function utilized in the SPIME v4 calibration pipeline:

def add_v4_physics_features(df):
    d = df.copy()
    temp, rh, pm25 = d['sample_temp'], d['sample_rh'], d['pm25']
    rh_clipped = np.clip(rh, 0.5, 99.8)
    rh_frac = rh_clipped / 100.0
    
    # Saturation Vapor Pressure e_s(T) in kPa (Tetens Formula)
    e_s = 0.61078 * np.exp((17.27 * temp) / (temp + 237.3))
    d['vpd'] = e_s * (1.0 - rh_frac)
    d['absolute_humidity'] = (216.7 * e_s * rh_frac) / (temp + 273.15)
    
    # Dew Point Temperature T_dew via Magnus Formula
    gamma = ((17.27 * temp) / (237.3 + temp)) + np.log(rh_frac)
    d['dew_point'] = (237.3 * gamma) / (17.27 - gamma)
    d['dew_point_depression'] = temp - d['dew_point']
    
    # Kohler Multi-Gamma Swelling Models
    for g in [0.05, 0.1, 0.15, 0.2, 0.25, 0.3, 0.4, 0.5]:
        d[f'kohler_gamma_{g}'] = pm25 * ((1.0 - rh_frac) ** g)
    d['rh_growth_ratio'] = rh_frac / (1.0 - rh_frac + 1e-4)
    d['pm25_kohler_scaled'] = pm25 / (1.0 + 0.15 * d['rh_growth_ratio'])
    d['rh_deliquescence_surge'] = np.where(rh_clipped > 75.0, np.exp((rh_clipped - 75.0) / 10.0), 0.0)
    
    # Multi-Window EWMA & Rolling Statistics
    for w in [1, 2, 3, 4, 6, 8, 12, 16, 24]:
        d[f'pm25_ewma_{w}'] = pm25.ewm(span=w, adjust=False).mean()
        d[f'pm25_roll_{w}_std'] = pm25.rolling(w, min_periods=1).std().fillna(0)
        d[f'pm25_roll_{w}_mean'] = pm25.rolling(w, min_periods=1).mean()
        d[f'pm25_roll_{w}_diff'] = pm25 - d[f'pm25_roll_{w}_mean']
        
    # Non-linear Cross Terms & Temporal Harmonics
    d['pm25_temp_interaction'] = pm25 * temp
    d['pm25_rh_interaction'] = pm25 * rh_clipped
    d['pm25_vpd_interaction'] = pm25 * d['vpd']
    d['temp_rh_ratio'] = temp / (rh_clipped + 1.0)
    d['log_pm25'] = np.log1p(np.maximum(pm25, 0))
    d['hour_sin'] = np.sin(2 * np.pi * d['hour_of_day'] / 24.0)
    d['hour_cos'] = np.cos(2 * np.pi * d['hour_of_day'] / 24.0)
    d['month_sin'] = np.sin(2 * np.pi * d['timestamp'].dt.month / 12.0)
    d['month_cos'] = np.cos(2 * np.pi * d['timestamp'].dt.month / 12.0)
    return d

APPENDIX B: HYPERPARAMETER TUNING CONFIGURATIONS

The following hyperparameter grid search spaces were evaluated across Time-Series Cross-Validation folds to optimize SPIME v4 base estimators:

Table B.1: Hyperparameter Search Spaces and Selected Optimal Parameter Configurations

| Model Algorithm | Evaluated Hyperparameter Grid | Selected Optimal Value |
| --- | --- | --- |
| Random Forest Regressor | n_estimators: [100, 300, 600] max_depth: [12, 16, 24, 32] max_features: [0.35, 0.55, 0.75, 1.0] | n_estimators: 600 max_depth: 32 max_features: 0.55 |
| ExtraTrees Regressor | n_estimators: [100, 300, 600] max_depth: [16, 24, 34] min_samples_split: [2, 4, 8] | n_estimators: 600 max_depth: 34 min_samples_split: 2 |
| LightGBM Regressor | num_leaves: [60, 120, 240, 380] learning_rate: [0.01, 0.015, 0.03] reg_alpha: [0.0, 0.005, 0.05] | num_leaves: 380 learning_rate: 0.015 reg_alpha: 0.005, reg_lambda: 0.05 |
| HistGradientBoosting | max_iter: [300, 600, 1200] learning_rate: [0.01, 0.015, 0.03] l2_regularization: [0.0, 0.005, 0.1] | max_iter: 1200 learning_rate: 0.015 l2_regularization: 0.005 |
| RidgeCV Meta-Learner | alphas: np.logspace(-4, 4, 30) fit_intercept: [True, False] | alpha: 0.0084 fit_intercept: True |