Chapter 5: Conclusion and Recommendations

This chapter concludes the study by summarizing the primary findings against the research objectives, detailing practical deployment recommendations for the African Sensor Evaluation and Training facility (Afri-SET), identifying study limitations, and suggesting avenues for future research.

5.1 Summary of Key Findings

This study developed, evaluated, and deployed a Stacked Physics-Informed Meta-Ensemble (SPIME v4) model to calibrate the low-cost QuantAQ Modulair-PM optical particle counter against a FEM Teledyne API T640 reference monitor at the University of Ghana, Legon, Accra.

The key findings corresponding to the research objectives are:

Failure of Factory Calibration: Uncalibrated factory sensor outputs yielded a negative R2 score (-0.0133) and an MAE of 5.30 μg/m3, demonstrating that European factory calibrations fail in West Africa's high-humidity and high-dust environment.

Impact of Physics-Informed Feature Engineering: Formulating 75 domain-specific features including multi-gamma Kohler hygroscopic swelling curves fRH,γ, Absolute Humidity (AH), Vapor Pressure Deficit (VPD), Dew Point Depression Tdew, and multi-window EWMA short-term trends provided explicit physical equations that eliminated reliance on spurious data correlations.

State-of-the-Art Calibration Accuracy: The SPIME v4 architecture achieved an R2 score of 0.9732 on test split data and 0.9632 on true unseen 6-month holdout data, reducing Mean Absolute Error (MAE) to 0.61 μg/m3 (a 66.3% error reduction over previous models).

Complete Elimination of Overfitting: By implementing 5-Fold out-of-fold cross-validation stacking with an L2-regularized RidgeCV meta-learner, the generalization gap was reduced to -0.0002, proving that the calibration model does not overfit to training set noise.

5.2 Practical Recommendations for Afri-SET and Network Operators

Based on the empirical results, the following practical recommendations are provided for deploying low-cost sensor networks across Sub-Saharan Africa:

Deploy SPIME v4 Calibration Pipelines: Network operators should replace simple linear or unregularized decision tree models with the SPIME v4 pipeline ([models/stacked_physics_meta_ensemble.joblib](file:///C:/Users/Hughe/Documents/SCHOOL/Final%20Year%20Project/Modulair-PM/models/stacked_physics_meta_ensemble.joblib)) to achieve reference-grade monitoring accuracy (R2≥0.96).

Mandate Meteorological Covariate Logging: Optical particle counters must always be deployed alongside high-precision relative humidity and temperature sensors. Calibration pipelines require ambient humidity and temperature to compute Kohler swelling and thermodynamic variables.

Incorporate Regional Harmattan Features: Calibration models deployed in West Africa must explicitly include Harmattan seasonal flags and dust storm indicators to prevent over-correcting dry mineral dust as humidity swell.

Periodic Re-Calibration Schedules: Co-located calibration models should be updated annually at hubs like Afri-SET to account for long-term optical diode aging and fan degradation.

5.3 Limitations of the Study

Single Site Spatial Scope: The co-located dataset was collected exclusively at the University of Ghana, Legon. While it captures severe seasonal variations, spatial transferability to other West African micro-environments (e.g. heavy industrial zones or traffic hotspots) requires further validation.

Focus on PM2.5: This study focused on fine particulate matter (PM2.5). Calibration of coarse particulates (PM10) and gaseous pollutants (NO2,O3,CO) was outside the scope.

5.4 Directions for Future Research

Spatial Transferability & Network Cross-Calibration: Evaluate SPIME v4 performance across multiple co-location sites in Sub-Saharan Africa to test spatial generalization.

Transfer Learning for Newly Deployed Units: Investigate transfer learning techniques to calibrate newly unboxed low-cost sensors using short 1-week co-location periods.

Integration into Real-Time IoT Web Portals: Embed the saved SPIME v4 pipeline into cloud-based REST API backend servers (prepare_backend.py) for automated real-time calibrated air quality map rendering.