# Chapter 1: Introduction

## 1.1 Background and Motivation

Ambient air pollution is one of the most pressing global environmental health crises of the twenty-first century, contributing to millions of premature deaths annually. Among various air pollutants, fine particulate matter with an aerodynamic diameter of 2.5 micrometers or less (PM2.5) is particularly hazardous. Because of its microscopic size, PM2.5 can penetrate deep into the human respiratory system, enter the bloodstream, and lead to severe cardiovascular and respiratory diseases, including lung cancer, ischemic heart disease, and stroke.

Despite the severe public health implications of air pollution, monitoring infrastructure remains critically sparse globally, with the most severe data gaps concentrated in low- and middle-income regions. Sub-Saharan Africa (SSA), in particular, suffers from a severe deficit of continuous, high-quality air quality monitoring. Establishing and maintaining traditional reference-grade monitoring stations—such as Beta Attenuation Monitors (BAM) or optical spectroscopy systems like the Teledyne T640—requires prohibitive capital costs (often exceeding $20,000 to $30,000 per unit), specialized technical expertise, stable power grids, and regular calibration maintenance. Consequently, vast urban populations across West Africa live in areas without any active ambient air quality monitoring.

In response to this monitoring deficit, low-cost sensors (LCS) have emerged as a disruptive and promising technology. These devices, which typically rely on optical particle counters (OPCs) to measure light scattering by ambient dust and smoke, are highly portable, consume minimal power, and cost a fraction of reference instruments (typically $200 to $2,000). By deploying dense networks of LCS, municipal authorities and researchers can map air pollution at high spatial and temporal resolutions, identifying local hot spots and tracking exposure levels in real-time.

However, the widespread deployment of LCS is hindered by significant data quality concerns. Optical low-cost sensors do not measure mass directly; instead, they count particles based on light scattering and convert these counts to mass concentrations using mathematical assumptions about particle density and refractive index. As a result, LCS readings are highly sensitive to:

Meteorological conditions: High relative humidity (RH) causes hygroscopic growth, where particles absorb water and swell, leading to systematic overestimation of particle size and mass.

Aerosol composition shifts: Seasonal wind patterns, particularly the dry, dust-laden Harmattan winds originating from the Sahara Desert, alter the physical and optical properties of ambient aerosols, leading to significant sensor drift and calibration mismatches.

Physical sensor degradation: Dust accumulation, laser aging, and fan degradation cause baseline drift over time.

To address these limitations, it is critical to develop robust, localized, and season-aware calibration models. Machine learning (ML) offers a powerful framework for this task, as it can model complex, non-linear interactions between raw sensor outputs, temperature, humidity, and seasonal factors (such as the Harmattan winds).

* * *

## 1.2 Air Quality Challenges in Sub-Saharan Africa and West Africa

The West African sub-region, particularly coastal cities like Accra, Ghana, faces a complex mixture of air pollution sources and meteorological drivers:

Anthropogenic Sources: Rapid urbanization and population growth have led to intensified vehicle emissions (often from older, poorly maintained diesel fleets), biomass burning for domestic cooking, open refuse burning, and industrial emissions.

Natural and Seasonal Sources: The region is heavily impacted by the Harmattan, a seasonal dry wind that blows from the Sahara Desert southwest across West Africa between November and March. The Harmattan carries vast quantities of mineral dust, which drastically increases ambient PM2.5 and PM10 concentrations.

Meteorological Extremes: Coastal West Africa experiences a tropical climate characterized by high year-round temperatures and alternating dry and wet monsoon seasons. High relative humidity during the wet season (often exceeding 80–90%) alters sensor responsiveness, while intense tropical rainfall events lead to atmospheric washout, rapidly scrubbing particulates from the boundary layer.

Standard factory calibrations for low-cost sensors, which are typically developed in temperate European or North American climates under low-humidity conditions with industrial combustion-dominated aerosols, fail catastrophically when deployed in West Africa. Under the high-humidity, high-dust conditions of Accra, uncalibrated LCS readings can deviate from reference instruments by over 100%.

* * *

## 1.3 The Afri-SET Initiative and the Modulair-PM Project

This research is situated within the active evaluation framework of the African Clean Air Technology Centre (Afri-SET), based at the Department of Physics, University of Ghana, Legon. Afri-SET acts as a regional hub for evaluating low-cost air quality sensors under local meteorological and environmental conditions in Sub-Saharan Africa. The facility co-locates candidate low-cost sensors with regulatory-grade reference instruments, creating co-located datasets that enable the training and testing of localized calibration models.

This project focuses on calibrating the QuantAQ Modulair-PM sensor (specifically unit MOD-PM-01068), a low-cost ambient air quality monitor that houses an optical particle counter (OPC-N3) alongside meteorological sensors for relative humidity, temperature, and atmospheric pressure. By co-locating this sensor at the University of Ghana campus with a US EPA Federal Equivalent Method (FEM) reference monitor—the Teledyne API T640—a high-resolution hourly dataset spanning the years 2024 to 2025 was established. This dataset captures both the wet monsoon season and the highly dusty dry Harmattan season, providing a unique foundation for developing advanced, meteorology-integrated calibration models.

* * *

## 1.4 Problem Statement

Low-cost air quality sensors (such as the QuantAQ Modulair-PM) offer a scalable solution to the air quality data desert in Accra, Ghana. However, their raw measurements suffer from severe inaccuracies due to environmental interference (hygroscopic particle growth under high humidity) and seasonal composition changes (coarse Saharan mineral dust during the Harmattan). Existing calibration methods, such as simple linear corrections, fail to capture the non-linear, interactive, and temporally dependent relationships between meteorological covariates and sensor errors. Furthermore, specific Saharan dust loadings are rarely integrated into low-cost sensor calibration pipelines in West Africa. Without robust, machine learning-driven calibration models that explicitly incorporate local meteorological contexts (like relative humidity and seasonal Harmattan winds), the data generated by low-cost sensors remains inadequate for regulatory compliance, epidemiological studies, or policy-making.

* * *

## 1.5 Research Aim and Objectives

### 1.5.1 Primary Aim

The primary aim of this study is to develop, evaluate, and deploy machine learning-based calibration models that improve the measurement accuracy of the QuantAQ Modulair-PM low-cost sensor for ambient PM2.5 prediction in Accra, Ghana, utilizing co-located reference data from a Teledyne T640 monitor.

### 1.5.2 Specific Objectives

Establish a Co-located Database: Align and preprocess hourly particulate and meteorological measurements from the QuantAQ Modulair-PM sensor and the Teledyne T640 reference monitor collected at the University of Ghana, Legon between 2024 and 2025.

Engineer Context-Aware Meteorological Features: Develop and integrate specialized features including relative humidity-temperature interaction terms, rolling statistics, temporal cyclic features, and a seasonal Harmattan score to account for regional climate patterns.

Train and Tune Machine Learning Classifiers: Build and optimize multiple calibration algorithms, including Linear Regression (baseline), Random Forest Regressor, and Multi-Layer Perceptron (MLP) Neural Networks, utilizing advanced hyperparameter search strategies.

Validate Generalization via Temporal Splitting: Evaluate the calibration models using chronological train-test splits and time-series cross-validation to prevent data leakage and simulate operational deployment.

Analyze Feature Importances and Residuals: Perform diagnostic error analysis, residual tests (normality and homoscedasticity), and interpret feature importances to understand the physical and environmental drivers of sensor error correction.

* * *

## 1.6 Scope of the Study

This study is limited to the calibration of the QuantAQ Modulair-PM sensor (unit MOD-PM-01068) co-located at the Department of Physics, University of Ghana, Legon, Accra. The analysis utilizes hourly data collected over the 2024–2025 period. Reference measurements are restricted to the co-located Teledyne API T640 Federal Equivalent Method (FEM) optical mass monitor. Environmental covariates include localized relative humidity, ambient temperature, barometric pressure, and a calculated regional Harmattan score. This study does not evaluate sensor spatial generalization to other geographic regions outside Accra, nor does it address calibration of gaseous pollutants (such as NOx or O3).

* * *

## 1.7 Structure of the Thesis

The remainder of this thesis is structured as follows:

Chapter 2: Literature Review conducts a detailed review of air pollution monitoring challenges, the working mechanisms of optical particle counters, environmental and meteorological interferences, existing machine learning calibration literature, and the unique atmospheric dynamics of the West African region.

Chapter 3: Methodology describes the experimental co-location setup, data preprocessing steps, feature engineering methods, the mathematical formulation of the machine learning algorithms (Random Forest and Neural Networks), the time-series validation strategy, and model evaluation metrics.

Chapter 4: Results and Discussion presents the model performances, comparative metrics, feature importance rankings, residual diagnostic plots, and a discussion of the models' behavior during the Harmattan and wet seasons.

Chapter 5: Conclusion and Recommendations summarizes the key findings, details the practical implementation considerations for Afri-SET, outlines the limitations of the study, and recommends directions for future research.

# Chapter 2: Literature Review

## 2.1 Fine Particulate Matter (PM2.5) and the Sub-Saharan African Context

Ambient particulate matter, specifically PM2.5 (particles with an aerodynamic diameter ≤2.5 μg/m3), is recognized as a major risk factor for global morbidity and mortality. Long-term exposure to elevated PM2.5 is epidemiologically linked to cardiovascular dysfunction, chronic obstructive pulmonary disease (COPD), respiratory infections, and systemic inflammation (Mauderly et al., 2008).

In Sub-Saharan Africa (SSA), the burden of air pollution-related disease is particularly severe due to the co-existence of heavy anthropogenic emissions (vehicle exhaust, domestic biomass burning, waste combustion) and natural aerosol sources. However, the region is characterized by an extreme "data desert" regarding air quality monitoring. Establishing a representative regulatory-grade monitoring network is often impossible for local governments due to high capital expenditures, complex maintenance schedules, stable energy requirements, and the necessity of highly trained personnel (DeSouza et al., 2017). Consequently, low-cost sensors (LCS) have been widely identified as the most viable mechanism to democratize air quality monitoring, provide localized spatial coverage, and protect public health.

* * *

## 2.2 Low-Cost Sensors vs. Reference Instruments

To evaluate the reliability of low-cost sensors, it is necessary to contrast their operating principles with reference-grade monitors.

### 2.2.1 Optical Particle Counters (OPCs) in LCS

Low-cost PM sensors, such as the QuantAQ Modulair-PM (which utilizes the Alphasense OPC-N3), typically operate using light scattering technologies. Inside the sensor, a sample of ambient air is drawn through a sensing chamber by a fan or pump. A laser beam illuminates the incoming stream of particles. When a particle intersects the laser beam, it scatters light, which is measured by a photo-diode detector positioned at a specific scattering angle.

The sensor counts individual light pulses (indicating particle count) and correlates the amplitude of the scattered light pulse with particle size based on Mie Scattering Theory. The particle counts are grouped into discrete size categories (known as "bins"). To convert these counts into mass concentration (μg/m3), the device applies the following mathematical conversion:

where Ni is the number concentration of bin i, Vi is the average volume of particles in bin i (under the assumption that all particles are perfectly spherical), and ρi is the assumed particle density.

This conversion process introduces systematic errors because:

Refractive Index Assumption: The intensity of scattered light depends heavily on the refractive index of the aerosol. LCS algorithms assume a constant refractive index, which varies significantly in real-world mixtures of black carbon, mineral dust, and organic aerosols.

Spherical Shape Assumption: Mineral dust particles are highly irregular and non-spherical, violating the spherical volume assumption and causing sizing errors.

Density Assumption: A uniform density is assumed (typically around 1.65 g/cm3), whereas carbonaceous soot (1.0--1.2 g/cm3) and Saharan mineral dust (2.5--2.7 g/cm3) have vastly different densities.

### 2.2.2 Reference-Grade Monitors: Teledyne API T640

In contrast, reference-grade instruments utilize more sophisticated, stable, and highly calibrated optical and physical methods. The Teledyne API T640 is a continuous particulate monitor that uses broad-spectrum white LED light-scattering spectroscopy to determine particulate mass concentrations. It measures light scattering over a wide range of angles and utilizes advanced refractive index compensation algorithms, enabling it to accurately size and count particles without the extreme sensitivity to chemical composition changes seen in low-cost sensors. The T640 is designated as a US EPA Federal Equivalent Method (FEM) for PM2.5 and PM10 monitoring, providing a highly reliable "ground truth" benchmark.

* * *

## 2.3 Environmental and Meteorological Interferences

Meteorological conditions exert a dominant influence on uncalibrated optical sensor readings, causing significant deviations from reference measurements.

### 2.3.1 Relative Humidity (RH) and Aerosol Hygroscopicity

The most significant meteorological driver of optical sensor error is relative humidity. Many atmospheric aerosols are hygroscopic, meaning they absorb moisture from the surrounding air. As relative humidity increases, particularly past the Deliquescence Relative Humidity (DRH) threshold (typically between 70% and 80%), particles undergo rapid hygroscopic growth. They absorb water, swell in size, and change their refractive index (approaching that of pure water, 1.33).

Because uncalibrated optical sensors cannot distinguish between a dry particulate and an water-swollen aerosol, they classify these grown particles into larger size bins and calculate a vastly inflated mass concentration. The relationship between dry particulate mass and relative humidity is often modeled using the f(RH) correction factor:

where κ represents the hygroscopic growth parameter of the aerosol mixture. In highly humid tropical environments like Accra, uncorrected hygroscopic growth is a major source of positive measurement bias.

### 2.3.2 Temperature and Volatilization

Ambient temperature affects both the physical state of the aerosol and the internal electronics of the sensor. High temperatures can cause the volatilization of semi-volatile organic compounds and ammonium nitrate from the aerosol phase to the gas phase, reducing the particulate mass measured by reference instruments but potentially leaving raw counts unchanged in optical sensors. Additionally, temperature fluctuations can affect laser diode stability and photo-detector response, necessitating temperature compensation in calibration models.

### 2.3.3 The Harmattan Wind and Saharan Mineral Dust

The dry winter season (November to March) in West Africa is dominated by the Harmattan winds. These trade winds transport vast quantities of coarse Saharan mineral dust across the sub-region. The Harmattan is characterized by:

Massive PM spikes: Ambient PM2.5 and PM10 concentrations routinely exceed 100 μg/m3 and can reach up to 500 μg/m3.

Low humidity: Relative humidity drops to extremely low levels (often <20%).

Coarse dust fraction: The aerosol distribution shifts toward larger, highly non-spherical, high-density mineral particles.

Under Harmattan conditions, optical particle counters face optical saturation and coincidence errors (where multiple particles pass through the laser beam simultaneously, registering as a single large particle). Furthermore, the high density (>2.5 g/cm3) and highly irregular shapes of dust particles violate the standard assumptions of the LCS mass conversion algorithm, causing uncalibrated sensors to underestimate the actual mass concentration. Thus, incorporating a "Harmattan Score" or seasonal indicator is essential for calibration models in West Africa.

* * *

## 2.4 Sensor Calibration Paradigms and Machine Learning

To correct these environmental biases, researchers co-locate low-cost sensors with reference instruments to build mathematical calibration models.

### 2.4.1 Classical Regression Models

Initial calibration attempts relied on Simple Linear Regression (SLR) or Multiple Linear Regression (MLR), expressing corrected PM as a linear combination of raw PM, temperature, and relative humidity:

While computationally simple, linear regression models fail to capture the non-linear, exponential behavior of hygroscopic particle growth at high relative humidity (>80%) and cannot model the complex interaction terms between temperature, humidity, wind patterns, and aerosol types.

### 2.4.2 Machine Learning Regressors

To overcome the limitations of linear models, recent literature has shifted toward non-linear machine learning architectures:

Random Forest (RF): Random Forest is an ensemble learning method that constructs a multitude of decision trees during training. It is highly robust, handles non-linear interactions natively, and is resilient to outliers. RF models have consistently outperformed linear models in sensor calibration studies, achieving high R2 scores by mapping environmental variables to sensor error corrections.

Multi-Layer Perceptron (MLP) Neural Networks: Artificial Neural Networks (ANNs) are capable of approximating any continuous non-linear function. By training multi-layer perceptrons with backpropagation, the model can automatically discover complex features and interaction terms from input variables (sensor PM, temperature, humidity, and temporal features). However, they require careful tuning of hidden layer sizes, activation functions, and regularization parameters to prevent overfitting.

Gradient Boosting (XGBoost/LightGBM): Sequential ensemble models that build decision trees iteratively to minimize residual error. They often yield marginal performance improvements over Random Forest but can be sensitive to hyperparameter configurations and missing values.

### 2.4.3 Validation Strategies and Data Leakage

In time-series calibration datasets, environmental conditions are highly autocorrelated (i.e., temperature at hour t is highly correlated with temperature at hour t+1). Standard random train-test splits (such as train_test_split with shuffle) lead to data leakage, where adjacent hours are split between the training and testing sets. This results in overly optimistic validation metrics that fail to reflect the model's actual performance when deployed to future unseen data. To ensure robust evaluation, recent research highlights the necessity of using chronological splits (e.g., training on 2024 data and testing on 2025 data) and time-series cross-validation, where the validation fold always succeeds the training fold in time.

# Chapter 3: Methodology

This chapter details the experimental design, data sources, preprocessing pipelines, feature engineering formulations, machine learning model architectures, and validation frameworks developed to calibrate the low-cost QuantAQ Modulair-PM sensor against the reference-grade Teledyne API T640 monitor.

* * *

## 3.1 Experimental Co-location and Data Sources

### 3.1.1 Co-location Site and Environmental Context

The experimental data collection was conducted at the African Clean Air Technology Centre (Afri-SET), situated within the Department of Physics at the University of Ghana, Legon, Accra (5.651221∘ N,−0.185531∘ W). Accra has a tropical savanna climate characterized by two distinct meteorological seasons:

The Wet Season (April to October): Influenced by the West African Monsoon, bringing frequent, heavy precipitation and high ambient relative humidity (often exceeding 80–90%).

The Dry Harmattan Season (November to March): Influenced by dry northeasterly trade winds blowing from the Sahara Desert. This season is characterized by heavy loading of coarse mineral dust, extremely low relative humidity (<30%), and elevated ambient particulate concentrations.

### 3.1.2 Low-Cost Sensor Unit: QuantAQ Modulair-PM

The low-cost sensor evaluated in this study is the QuantAQ Modulair-PM (specifically unit MOD-PM-01068). This device is an integrated air quality monitor containing:

Alphasense OPC-N3 Optical Particle Counter: Measures particle counts across 24 size bins spanning 0.35 μm to 40 μm. It outputs raw particle counts, calculated PM mass concentrations (pm1, pm25, pm10), and diagnostic parameters (opcn3_temp, opcn3_rh).

Meteorological Sensors: Measures ambient temperature (sample_temp in ∘C), relative humidity (sample_rh as a percentage), and barometric pressure (sample_pres in hPa).

QuantAQ Calibration Algorithm: Outputs proprietary corrected PM values (pm25_env, pm10_env). In this study, we utilize the raw sensor reading (pm25) as our primary feature to build independent, localized calibration models.

### 3.1.3 Reference Monitor: Teledyne API T640

The ground truth reference data was collected using a co-located Teledyne API T640 particulate monitor. The T640 utilizes broad-spectrum LED light-scattering spectroscopy to provide continuous, high-resolution hourly measurements of PM2.5 mass concentrations. It is a US EPA Federal Equivalent Method (FEM) instrument with an operating range of 0 to 10,000 μg/m3 and a measurement resolution of <0.1 μg/m3.

* * *

## 3.2 Data Preprocessing and Alignment

### 3.2.1 Temporal Alignment and Reindexing

The low-cost sensor collects data at a high temporal resolution (typically 1-minute intervals), whereas the Teledyne T640 reference monitor logs hourly averages. To align the datasets:

The 1-minute Modulair-PM sensor data was aggregated to hourly averages using the mean.

Both datasets were converted to Coordinated Universal Time (UTC) and aligned by index.

Due to potential offsets and minor gaps, the sensor data was reindexed to match the reference database index (Timestamp (UTC+0)), ensuring complete temporal alignment.

### 3.2.2 Quality Control and Outlier Handling

To ensure data integrity, data quality filters were applied:

Negative Values: Reference PM2.5 measurements below 0 μg/m3 (due to instrument zero-drift) were flagged and removed.

Instrument Malfunction Flags: Data points where the Modulair-PM sensor flagged device status errors (flag != 0) were excluded.

Missing Data: Gaps of ≤2 consecutive hours in the low-cost sensor data were filled using linear interpolation. Gaps exceeding 2 hours were dropped from the training and testing sets to prevent introducing artificial signals.

Outlier Detection: The Interquartile Range (IQR) method was applied to raw particulate measurements. Extreme spikes exceeding Q3+3×IQR were cross-checked with reference measurements; if the reference monitor did not register a corresponding spike, the sensor data was treated as an instrumental error and removed. Genuine high-pollution events (such as dust storms) where both instruments spiked were preserved.

* * *

## 3.3 Feature Engineering

To capture the physical and chemical behavior of the sensor and its environmental interferences, several domain-specific features were engineered.

### 3.3.1 Aerosol Hygroscopicity and Temperature Interaction

To model the hygroscopic growth of particles in high relative humidity, we engineered an interaction term between the raw sensor measurement and relative humidity:

This interaction allows the model to scale its correction non-linearly when high relative humidity and high particulate loadings occur simultaneously.

### 3.3.2 Temporal and Diurnal Features

Air pollution concentrations and boundary layer heights follow strong diurnal and weekly cycles. These were captured using:

Hour of Day (H∈[0,23]): Represents the diurnal traffic, boundary layer height, and temperature variations.

Day of Week (D∈[0,6]): Captures anthropogenic activity differences between weekdays and weekends.

### 3.3.3 Rolling Temporal Statistics

To incorporate atmospheric continuity and local trends, short-term rolling statistics were calculated on the raw sensor measurements using a rolling window (w=3 hours):

Rolling Mean: Captures the short-term background trend:

Rolling Standard Deviation: Captures short-term volatility and rapid changes in concentration:

To avoid missing values at the beginning of the time-series, the first w−1 values were backfilled with the raw pm25 value.

### 3.3.4 Harmattan and Dust Storm Indicators

To handle the unique dust loading during the dry winter months, three features were developed:

is_harmattan_month: A binary flag (1 if the month is in November, December, January, February, or March; 0 otherwise).

is_harmattan: A localized seasonal indicator mapping active dust transport periods based on low humidity and high coarse-particle ratios.

harmattan_score: A continuous variable estimating dust severity:

\text{harmattan\_score} = \begin{cases}

f(\text{coarse\_ratio}, \text{RH}) & \text{if dry season months} \\

0 & \text{otherwise}

\end{cases}

### 3.3.5 Feature Scaling

For algorithms sensitive to scale (such as Neural Networks), continuous features were standardized to have zero mean and unit variance using the StandardScaler:

where μ and σ are the mean and standard deviation of the feature calculated only from the training partition to prevent data leakage.

* * *

## 3.4 Model Architectures and Training

Three modeling paradigms were implemented and compared:

### 3.4.1 Baseline Model: Linear Regression

A standard Multiple Linear Regression (MLR) model was implemented as a baseline:

### 3.4.2 Ensemble Model: Random Forest Regressor

The primary machine learning model is a Random Forest Regressor, which constructs an ensemble of decision trees using bootstrap aggregation (bagging) and random feature selection:

Structure: Trained with 200 decision trees (n_estimators=200).

Hyperparameter Optimization: Conducted using Grid Search (GridSearchCV) over a hyperparameter space:

n_estimators: [200,500]

max_depth: [10,15,20,None]

min_samples_split: [2,5,10]

Objective Function: Minimize the Mean Absolute Error (MAE) during training.

### 3.4.3 Deep Learning Model: Multi-Layer Perceptron (MLP) Regressor

To capture highly non-linear high-dimensional interactions, an artificial neural network was built using Scikit-Learn's MLPRegressor:

Architecture: A feedforward neural network comprising:

Input layer mapping the engineered feature vector.

A hidden layer containing 128 neurons (hidden_layer_sizes=(128,)).

Output layer containing a single neuron with linear activation to output the calibrated PM2.5 concentration.

Hyperparameters:

Activation function: Hyperbolic tangent (activation='tanh').

Optimization solver: Adam (solver='adam').

Regularization parameter: L2 penalty parameter α=0.01 (alpha=0.01).

Batch size: 64 (batch_size=64).

Training control: Early stopping enabled (early_stopping=True) to prevent overfitting.

* * *

## 3.5 Model Validation and Testing Strategy

To evaluate how well the models generalize to unseen future data and prevent temporal data leakage:

Chronological Train-Test Split: The co-located dataset was split chronologically rather than randomly. A training partition (75% or 80%) consisting of earlier records (primarily 2024 data) was used to train the models, while a holdout test partition (20% to 25%) consisting of subsequent records (2025 data) was used to evaluate generalization.

Time-Series Cross-Validation: A 5-fold Time Series Split (TimeSeriesSplit) was applied during hyperparameter optimization, where the training set in each fold grows sequentially, and the validation set is always chronologically subsequent to the training set. This simulates a real-world deployment scenario where models calibrated on past data must predict future air quality.

* * *

## 3.6 Evaluation Framework

Model performance is assessed quantitatively using three primary metrics:

Coefficient of Determination (R2 Score): Measures the proportion of variance in the reference data explained by the calibration model:

Mean Absolute Error (MAE): Measures the average magnitude of absolute errors (expressed in μg/m3):

Root Mean Squared Error (RMSE): Penalizes larger deviations, making it sensitive to extreme prediction errors:

In these equations, yi is the actual reference PM2.5 value from the Teledyne T640, y^i is the calibrated value predicted by the model, y is the mean of the actual reference values, and n is the total number of test samples.

In addition to quantitative metrics, Residual Diagnostics were conducted:

Residual Distribution Plotting: Histograms and Kernel Density Estimation (KDE) to verify if prediction errors are normally distributed.

Quantile-Quantile (Q-Q) Plots: Standardized residuals plotted against theoretical normal quantiles to assess deviations from normality.

Heteroscedasticity Analysis: Scatter plots of residuals versus predicted values to determine if the error variance is constant across the range of predictions.

Feature Importance Analysis: Extraction of tree-based feature importances from the Random Forest model to identify the physical parameters most critical to correcting sensor readings.

Mass=iNi×Vi×ρi

PMobserved=PMdry×f(RH)

f(RH)=1+κRH100−RH

PMcalibrated=β0+β1PMraw+β2T+β3RH

InteractionPM-RH=pm25×sample\_rh

rolling\_mean\_pm25t=1wi=0w−1pm25t−i

rolling\_std\_pm25t=1w−1i=0w−1(pm25t−i−rolling\_mean\_pm25t)2

z=x−μσ

PM25reference=β0+β1pm25raw+β2sample\_temp+β3sample\_rh+β4InteractionPM-RH+ϵ

R2=1−i=1n(yi−y^i)2i=1n(yi−y)2

MAE=1ni=1n|yi−y^i|

RMSE=1ni=1n(yi−y^i)2